| LOT | ITEM | QUESTION | ANSWER |
| --- | --- | --- | --- |
| 8 |  |  |  |
|  | 341 | Electronic Force Transducers |  |
|  |  | For compression loads only, or for compression and tension loads? | Compression and tension loads for all transducers to 1000 kN; compression only above that. |
|  |  | Which class according to ISO 376 is required? | ISO 376 Class 00 from 20% to 100% of range for all transducers |
|  |  | Do you need a measuring amplifier for each force transducer or should one measuring amplifier be used with all force transducers? | One dual input measuring amplifier for all transducers together |
|  | 342 | Universal Force Machine up to 6000kN |  |
|  |  | Are you looking for a fully automatic hydraulic calibration machine or for a manually operated hydraulic machine with frame and hand pump? | Fully automatic machine. |
|  |  | Must the machine only be for compression loads? | Must match Item 341, so both compression and tension loads for all transducers to 1000 kN and compression only above that. |
|  |  | Is a 5000 kN machine OK instead of a 6000 kN machine? | The machine should match the range of transducers of Item 341, but these transducers can handle forces beyond the specified range. Therefore, 6000 kN is preferred, but 5000 kN may also be acceptable and can be considered compliant. |
|  | 343 | Torque Transducers |  |
|  |  | Is class 0.05 according to ISO 51309 required? | Yes |
|  |  | Do you need a measuring amplifier for each torque transducer or should one measuring amplifier be used with all torque transducers? | One dual input amplifier for all of them together. |
|  | 414 | Torque Generator |  |
|  |  | Are you looking for a fully automatic machine like a jockey weight machine or for a hand operated torque tester? | No jockey weight machine. Hand operated or motor operated are both acceptable, as long as maximum specified torque and stability can be realised with ease. |
|  | 380 | Electronic Force Transducers |  |
|  |  | For compression loads only, or for compression and tension loads? | Compression and tension loads for all transducers to 1000 kN; compression only above that. |
|  |  | Which class according to ISO 376 is required? | ISO 376 Class 0.5 from 20% to 100% of range for all transducers |
|  |  | Do you need a measuring amplifier for each force transducer or should one measuring amplifier be used with all force transducers? | One dual input amplifier for all of them together. |