# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

Contract title: Supply of laboratory and metrology equipment for testing and calibration laboratories in Ghana

# Publication reference: EuropeAid/133510/D/SUP/GH

Column 1-2 should be completed by the Contracting Authority

Column 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

LOT 10: METROLOGY EQUIPMENT FOR VOLUME

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 368 | Multi-channel pipette calibrator (12 channels) + controller hardware and software (calculation unit) - Features: - Gravimetric calibration in compliance with DIN EN ISO 8655 - Test volumes of water are pipetted simultaneously into up to twelve weighing tubes located next to one other - Automatic weighing of weighing tubes in parallel by a 12-channel balance - Integrated evaporation trap ensuring values of the highest precision - Tube drainage by suction device (no need to remove the tubes) - Weigh cells are tared automatically after the measured values have been recorded - Computer records the weight measurements, calculates the pipetted volumes and immediately generates a report containing individual and mean values as well as the standard deviation - Integrated communications port, adjustable to specific customer needs - Specifications: - Time required to test a 12-channel pipette (= 360 individual measurements) is about 10 minutes (operating time of the test unit and manual pipetting) - Nominal volume of multichannel piston-operated pipettes: 10 μl to 1.2 ml - Resolution 0.01 mg - Tare range 12 g - Repeatability (standard deviation): < 0.02 mg - Uncertainty (k=2): < 0.04 μl (in the range of 0 to 1,200 μl) - Response time (average) Approx. 10 s - Sensitivity to temperature drift between +10 °C and 30 °C: < ± 1x10 ppm/K - Including: - Control/data interface, cable to PC - System controller PC with control/analysis software for generating calibration certificates - calibration certificate from an accredited calibration laboratory - User manuals in English - Service/repair contract and spares for 3 years |  |  |  |
| 370 | Fuel dispenser check measures set 5, 10 and 20 L incl. carrying cases - Features: - Stainless steel check measures, constructed according to the international recommendation OIML R-120 - Specifications: - Scale length: -3 % to +3 % from the nominal value, scale indicating deviation from nominal in ml - Maximum deviation after adjustment of the scale zero to the nominal volume: 0.05 % - Scale adjustment mechanism must be sealable to prevent inadvertent adjustment - Calibration uncertainty: 0.05 % - Including: - calibration certificate from an accredited calibration laboratory - Carrying case |  |  |  |
| 376 | Set of measuring cylinders (graduated), 100 ml, 200 ml, 500 ml, 1 L and 2 L, grade A - Features: - Boron silicate glass or similar - Cylindrical shape on large base - Graduations, main: 10 %, intermediate: 5 %, small: 1 % of nominal volume - Graduations from 10 % to 100 % of nominal volume - Specifications: - Accuracy grade: A, or 0.5 % maximum deviation of nominal volume - Including: - Protective carrying case for storage and transport of cylinders - calibration certificate from an accredited calibration laboratory |  |  |  |
| 378 | Water meter test bench DN15 to DN40 (upgrade of existing system), with controller hardware and software - Description: - Existing system at GSA includes a water storage tank, pump, reference meter, pipe works and valves - The upgrade includes: - Additional valves and pipe work to be connected to the existing system - Adaptors, valves, pipe work and connectors/adaptors for connection in series of at least 8 water meters of various pipe diameters and connection standards - A 500 L tank, mounted on a 600 kg precision scale, with non-touching water inlet spout and non touching draining valve, for the calibration of water meters by gravimetric method - Outlet from gravimetric tank into a drain tank with pump to recycle the water into the storage tank - Existing pumping capacity, 300 L/min may be used to realise selectable flow rates from minimum 5L/min to maximum 300 L/min - Automated start/stop control, flow control, measurement data storage and analysis - Existing flow meter (ultrasonic 30 to 3000 L/min, 0.15%) can be used for monitoring flow rates - Specifications: - System to accommodate volumetric and velocity type water meters with flow rates from 5 L/min to 300 L/min - Calibration uncertainty to be achieved: 0.05 % - Maximum length of meters under test: 300 mm - Resolution of 600 kg precision scale: 1 g to 100 kg, 10 g above 100 kg - Maximum deviation (including uncertainty) of precision scale: 0.025 % of reading from 20 kg to 600 kg - Including: - System controller PC and software for mass readings - Facility for measuring water temperature at 0.1 °C uncertainty with data communication to PC - Installation and test report in English - calibration certificate from an accredited calibration laboratory - Operation manual in English - Service/repair contract and spares for 3 years |  |  |  |

Requested training:

| Spec | Equipment | Training |
| --- | --- | --- |
| 368 | Multi-channel pipette calibrator | ½ day for max. 2 persons |
| 378 | Water meter test bench | 1 day for max. 4 persons |

General Remarks:

- If calibration is required, the calibration certificate must be issued by a calibration laboratory accredited by an accreditation body, which is a signatory of the Mutual Recognition Arrangement of the International Laboratory Accreditation Cooperation (ILAC-MRA).

- All electrical equipment should be suitable for operating on 220-240V / 50-60 Hz (unless otherwise specified), and provided with a BS-1363 plug, or with a suitable adapter.

- Software that is required for the realisation of functionalities of equipment must be the latest version, and should be compatible with a Windows 8 or equivalent operating system. Software updates should be provided for free during at least for a period of 36 months.

- For computers/system controllers and related supplementary equipment (e.g. flat screen, printer, UPS) that are considered to be an integrated part of the equipment configuration and that are not further specified, the contractor will provide systems of good quality and reputable brands, based on latest technology, and latest software for operating systems (Windows 8, or equivalent), for office applications (MS Office, or equivalent), and anti-virus protection at least for a period of 36 months.

Additional information on selected items for better understanding the nature of the repairs / refurbishment required

| Spec | Equipment | Description |
| --- | --- | --- |
| 378 | Water meter test bench DN15 to DN40 (upgrade of existing system), with controller hardware and software | The specified calibration uncertainty of 0.05 % is the uncertainty at k=2 level of the master meter calibration with the gravimetric method, using the 500 L tank and scale. This uncertainty should be achieved for pumped volumes of 100 L to 500 L. This system shall be used to calibrate/standardise the master meter, which will subsequently be used to calibrate the meters under test. The uncertainty of the test rig as a whole will be dependent on the performance of the existing storage tanks, pipes, valves and reference meter assembly and is therefore not specified in the tender. A site visit will be arranged on 15th April 2013 for tenderers to establish any further details about the existing equipment and laboratory lay-out, but attendance is not obligatory. Guidance on the construction and capabilities of the test bench - Environmental conditions to be assumed: - The laboratory is air-conditioned. - Ambient temperature and humidity in the lab are monitored. - Water temperature (outdoor tank): 20 C to 30 C - Ambient temperature range (in the lab): 20 C to 26 C - Ambient relative humidity range: 45 % to 95 % - Ambient atmospheric pressure: 86 kPa to 108 kPa - During each test, the temperature and relative humidity in the lab is expected to vary by not more than 3 C or 10% respectively. - The maximum uncertainty in the measurement of ambient temperature will not exceed 1 C. - Existing equipment: - Flow test rig for various types of flow meters 2” to 6” diameters, temperature compensated system with automatic batch controller. - Reference standard master meter, ultrasonic 30 to 3000 L/min, 0.15%, traceable to UKAS. - Water Meter Test Bench operational parameters to be achieved: - A total of 8 meters under test should be able to be fitted in series in horizontal position. - The test bench should allow any number of meters up to the maximum to be fitted in the system and be tested. - The construction of the test bench should allow at least two meters under test to be mounted in vertical position in series. - Maximum inlet pressure as a function of flow rate is determined by the existing inlet pump and assembly and may be assumed to be able to reach at least 10 bar. - Outlet pressure of any meter under test should not be less than 0.03 MPa (0.3 bar). - The pressure upstream of the meter should not vary by more than 10 %. The performance is given by the existing pump and control system. - The maximum uncertainty in the measurement of pressure shall not be more than 5 % of the measured value. - The calibration rig should be built to an operating temperature of up to 30°C. - Meter sizes: G¾B, G1B, G1¼B, G1½B, G2B, DN15, DN20, DN 25, DN 32, DN40. - Max. flow rate for water meters: 300 L/min. - Min. flow rate 5 L/min. - Uncertainty: - Using the standardisation of the reference meter by the gravimetric method, a calibration uncertainty of 0.1 % (k=2) of the meters under test is expected. - Operation: - It shall be possible to operate all relevant pumps and valves on the supply- and reference unit from the control interface, both as automatic calibration and manual mode. - Choice of flow direction, return to reservoir or to main reference shall be operated from the control interface. - Flow rate shall be controlled from the control interface. - Calibration time shall be controlled from the control interface. - Measurement procedure: - At start of a calibration, a flow rate or flow sequence shall be chosen. The system then chooses pump and meter settings to use to get the best performance at the lowest uncertainty. - It must be possible to choose whether to run against the main reference (weighing) system or not. - Using reference meters it shall be with standing start/stop. - Using the main reference (for calibration of reference meters) the system must be designed to operate with dynamic measuring. It shall be possible to operate for a chosen time, a chosen volume. - Parameters of influence: - Fluid temperature measuring: - in reservoir - before the reference meters - after test string - Pressure measuring: - before reference meters - before and after test string - Control panel: - The Control interface shall be a system overview either as a panel or as a screen picture. - Main reference: - Balance with a resolution of 1g. - Divider with feedback to the controls for start/stop of measuring counters. - Suggested materials: - Pipes and fittings: sanitary components in AISI 316 quality (ISO) - Electro-pneumatics valves: butterfly valves in AISI 316 quality (ISO) - Sheets and profiles: AISI 316 quality (ISO) - Connection to supply: - The calibration rig is fitted with connectors, on the respective unit for main power, compressed air, water supply and waste drain. - Valves: - The supply- and reference section shall include valves mounted next to the pumps and the reference meters. - Blocking valves are placed before and after the test string. - All relevant valves shall be PN10 butterfly valves and it shall be possible to operate them from the user-interface. - Other types of valves can be used for bleed. - Bleed valves for evacuation of air must not be active under calibration. - Type (general): Butterfly - PN: 10 - Temperature sensors: - All sensors shall be Pt100, EN 60 751 - All sensors shall be readable from the user-interface. - Type: Pt100 (EN 60 751) - Class: IP 65 - Min. measuring range: 0 – 35 °C - Resolution on display: 0.1 °C - Pressure sensors: - All sensors shall be according to DIN 16 086 - Read-out in bar or kPa - All sensors shall be readable from the user-interface. - Read-out: Bar or kPa - Type: DIN 16 086 - Resolution range: 0.01 bar or 0.1 kPa - Weights: - Weights, in pieces for calibration of the balance are available at the institute and need not be included in the tender. - Weighing tank: - The main reference shall include a symmetrical 500 litres of water weighing tank with a draining valve. - Divider: - The main reference shall include a divider. - The switching time of the divider shall be less than 0.1 s. - No leak or splash of liquid must occur. - Control system: - The control system shall have a user-interface with an overview of the system, shown as a picture on the screen or on the control panel. - From this interface it should be able to operate all pumps and valves. - Readings of temperature, flow, volume, pressure and weight shall be possible. - Function: - The Control system shall have 3 possible modes: Manual It shall be possible to operate all pumps and valves. It shall be possible to read flow, pressure, temperature and weight. Automatic: The system shall be able to choose pump and reference meter settings after selected flow / time. There must be a sequence for evacuation of air where it shall be possible to manual start / stop a low flow. It shall be possible to read flow, pressure, temperature and weight. The flow shall remain until a preset volume, a preset time or until the manual stop. After stop the volume shall be shown at a resolution not less than or equal to the requirements for the reference meters. Calibration of reference meter: The Control system shall be able to tare and read the weight from the reference scale. The Control system shall be able to calculate the correction of the reference meters. It shall be possible to read flow, pressure, temperature and weight. - Pumps, reference meter and meters under test shall be protected against too large flow and pressure. - The safety system shall also protect against overload and shall ensure that the system not will be able to run dry. - Reporting: - A report showing the various calibration data and conditions shall be created. - Volume measured by the master meter - Flow rate - Temperatures - Pressures - Weight - Correction of master meter - In case of master meter calibration, additional parameters to be reported: - Calculated meter correction - Previous correction - The difference between current correction and the previous correction in percent and there shall be a possibility to have accept levels for the correction. - Operator shall have a possibility to choose the points of calibration and volumes for the reference meter. - System shall be able to calibrate and report for 7 different flow rates. - Suggested system set-up: weighing tank |