# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

Contract title: Supply of laboratory and metrology equipment for testing and calibration laboratories in Ghana

# Publication reference: EuropeAid/133510/D/SUP/GH

Column 1-2 should be completed by the Contracting Authority

Column 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

LOT 11: METROLOGY EQUIPMENT FOR ACOUSTICS AND VIBRATION

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 382 | Microphone and sound level meter calibration system, including microphones and controller hardware and software Sound Level Meter (SLM) calibration and microphone calibration system with advanced sound level meter/analyser - Features: - SLM calibration system consisting of at least: - Data analyser with a frequency range from DC – 50 KHz - Signal generator with sine, random and burst tones required to perform calibration of sound level meter according to IEC standard - Multitone acoustic calibrator with frequency range from 31.5 Hz to 12,5 KHz - Low noise microphone with accessories to perform noise check according to standard - Microphone calibration system, consisting of at least: - Microphone Calibration Module - Comparison Calibration Coupler (WS2) with adaptors for WS1 and WS3 - Microphone and Coupler Stand - Electrostatic Actuators with adapters for WS3 calibration - Reference microphone with matching pre-amplifier - Working microphone with matching pre-amplifier - Specifications for the SLM calibration system: - Trial signals according to sound level meter IEC 60651, IEC 60804 and IEC 61672 standards - Acoustical testing on the frequency range 31,5 Hz to 12,5 kHz with an acoustical calibrator - Electrical testing (simulation) - Operation modes: automatic, semi-automatic and manual - Software for testing all types of SLM according to the above standard for at least the following: - Absolute Acoustical Sensitivity Level - RMS detectors (Crest factor 3, 5, and 10) for different types of signals: sinus, rectangular pulses - Generation of sinus signals of 20 Hz to 12,5 kHz and Burst pulses - Determining Electrical Level for Lref at 1 kHz - Electrical Inherent Noise Level (A, B, C, Lin) - Frequency Response (A, B, C, Lin) - Linearity Range (SPL, Leq and SEL) - Level Range Control - Time Weighting - Response to Single Burst (Fast, Slow, Impulse) - Response to a Continuous Sequence of Bursts (Impulse) - Peak - RMS Detector (Crest factor 3, 5 and 10) - Sine Bursts - Rectangular Pulses - Overload Indication - Sine Signals, Inverse A - Time averaging - Pulse Range - Specifications for the microphone calibration system: - Calibration of Working Standard Microphones (IEC61094-4, LS1, LS2 and LS3) by using the methods described in IEC 61094-5 (sensitivity comparison) and IEC 61094-6 (frequency response by electrostatic actuator) - Sensitivity Measurement Frequency: 250 Hz or 1000 Hz - Measurement Range of Frequency Response: 20 Hz to 50 kHz (response normalized at Sensitivity Measurement Frequency) - Corrections to Actuator Response: Free-field, Diffuse-field, Pressure-field (Depending on type of microphone) - Uncertainty (k=2) of Sensitivity Calibrations: 0.07 dB – 0.14 dB (depending on type of microphone) - Uncertainty (k=2) of Actuator Response Calibration (ref. 250 Hz): - < 0.10 dB at 20 Hz - < 0.10 dB at 5 kHz - < 0.12 dB at 10 kHz - < 0.20 dB at 40 kHz - < 0.30 dB at 100 kHz - Including: - Computer, monitor and printer as controller for the system, including the required software package - Sound level meter calibration and microphone calibration software including database functionality for storage and reporting of calibration results - All cables, adaptors and accessories necessary for proper operation and all equipment are operated by power supply 230VAC and/or battery - Calibration certificates from an accredited laboratory for all reference equipment - Complete technical documentation in English including operating, service and reference manuals - Delivery and on-site installation, testing and commissioning - Service/repair contract and spares for 3 years |  |  |  |
| 383 | Sound calibrator 1 kHz, 94 and 114 dB - Features: - Multifunction Acoustic Calibrator - For calibration and comprehensive check of ½” and 1/4” microphone set-ups and acoustic instruments conforming to IEC651 and ANSI S 1.4-1983 such as sound level meters, dose meters, etc. - Compliant with IEC60942 class 1 - Specifications: - Calibration levels of 94 dB, 104 dB and 114 dB with an uncertainty in the SPL of less than 0,2 dB for SPL at 94 dB with wide frequency range: tones from 31.5 Hz to 16 kHz in octave steps, plus tone at 12.5 kHz - Including: - Calibration certificates from an accredited laboratory for all reference equipment - Complete technical documentation in English including operating, service and reference manuals - Extended warranty plus fast delivery of spare parts |  |  |  |
| 384 | Sound Level Meter - Features: - Free field microphone - Real-time analysis of sound in 1/1- and 1/3-octave bands - Event detection - Loudness and noise rating measurements - Frequency weightings A, B, C, Z (linear) and time weightings F, S, I - 3 Hz – 20 kHz broadband linear frequency range - Ldn, Lden, Lday, Levening, Lnight parameters directly on the instrument readings - Analysis of time histories for broadband parameters and spectra (Logging) - Automatic detection of and correction for windscreen - Robust and environmentally protected for continuous outdoor use - Documentation of measurements using text, voice and metadata annotations - Documentation of measurements through recording of measured signals - Logging up to 10 broadband parameters and 3 spectral parameters - Large high-resolution screen - Data storage - Communication via USB, LAN, or GPRS/3G modems - Timers for automatic start of measurement - Specifications: - Class 1 sound measurements to the latest international standards - Sound recording in 24 bit resolution - Tone assessment using 1/3-octave methods - Dynamic range in excess of 120 dB - Including: - PC software for archiving, previewing and exporting data; software maintenance and remote online display - Rechargeable battery and battery charger - Calibration certificates from an accredited laboratory for all reference equipment - Complete technical documentation in English including operating, service and reference manuals - Extended warranty plus fast delivery of spare parts |  |  |  |

Requested training:

| Spec | Equipment | Training |
| --- | --- | --- |
| 382 383 384 | Microphone and sound level calibration system Sound calibrator Sound level meter | 5 days (for all specs combined) for max. 4 persons |

General Remarks:

- If calibration is required, the calibration certificate must be issued by a calibration laboratory accredited by an accreditation body, which is a signatory of the Mutual Recognition Arrangement of the International Laboratory Accreditation Cooperation (ILAC-MRA).

- All electrical equipment should be suitable for operating on 220-240V / 50-60 Hz (unless otherwise specified), and provided with a BS-1363 plug, or with a suitable adapter.

- Software that is required for the realisation of functionalities of equipment must be the latest version, and should be compatible with a Windows 8 or equivalent operating system. Software updates should be provided for free during at least for a period of 36 months.

- For computers/system controllers and related supplementary equipment (e.g. flat screen, printer, UPS) that are considered to be an integrated part of the equipment configuration and that are not further specified, the contractor will provide systems of good quality and reputable brands, based on latest technology, and latest software for operating systems (Windows 8, or equivalent), for office applications (MS Office, or equivalent), and anti-virus protection at least for a period of 36 months.