# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

Contract title: Supply of laboratory and metrology equipment for testing and calibration laboratories in Ghana

# Publication reference: EuropeAid/133510/D/SUP/GH

Column 1-2 should be completed by the Contracting Authority

Column 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

LOT 2: CHROMATOGRAPHY

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 008 | Nitrogen Generator for GC 1. Flow rate, L/min: 3.0 2. Purity: <10 ppm O2 3. Electrical supply: 240V, 50Hz 4. Connection outlet: ⅛ inch Swagelok 5. With an in-built compressor 6. Digital hours counter |  |  |  |
| 009 | HPLC MS/MS A complete system comprising an auto-sampler, quaternary gradient solvent delivery system and an MS/MS Detector - supplied with a PC based system controller and Windows 8 or equivalent or equivalent compatible software to control the operation of the liquid chromatograph, mass spectrometer and for the acquisition and processing of data, together with a laser printer. - Solvent Delivery System: - Quaternary pump capable of mixing up to 4 components of an eluting mixture - Flow range, capable of being set to deliver a flow rate of ≤ 10 to ≥ 2000µl/min with ≤ 1.0 µml/min increments - Flow precision at least 0.15% RSD - Composition precision least 0.50% SD - Variable stroke volume - Maximum operating pressure at least 1000 bar - Complete with appropriate on-line vacuum degasser - Autosampler- Programmable - Fitted with Rheodyne type injector with 100 µl sample loop - Injected volume – 0.1 to 10 µl - ability to extend up to 250 µl - Precision ≤ 0.5% RSD from 0.2 – 2 µl, ≤ 1% from 2 – 10 µl - Linearity >0.999 coefficient of deviation at 0.2 - 10 µl - Sample capacity: at least 100 sample vials (2 ml) with random access to any vial - MS Quadrupole detector - Capable of operating in the MS – MS mode - Analyser mass range: 2 to > 2000u - Scan rates: ≤ 500 to ≥ 10000 u/sec - Resolution: ≤ unit mass over the entire range - Mass axis stability: ≤ ± 0.1 u over 24 hours - Manifold temperature: independent control; ≤ 25 °C to ≥ 70 °C - Ion detector: positive or negative ions, constant post acceleration voltage - ESI must be provided as standard and a combined ESI/APCI must be provided - Turbo-molecular pump: dual stage - Interface for connection to outlet of HPLC column and for direct injection of solutions into the mass detector - Chromatography & Mass spectrometry Software - Windows 8 or equivalent or equivalent compatible software for data acquisition and processing, and for control of all instrument parameters via a PC - Quantification and Calibration possibilities - Quality assurance system - Data import and export to other software - Possibility to upgrade the system for new instruments - A desk top computer suitable for running the software, equipped with Windows 8 or equivalent, high resolution LCD or LED monitor, and a laser printer (printer to be supplied with high volume cartridges sufficient for approx. 8000 to 10.000 pages). - Service/repair contract and spares for 3 years |  |  |  |
| 010 | Gas Chromatograph – Triple Quadrupole Mass Spectrometer Complete with auto-sampler, fitted with a split/splitless capillary injector and a mass selective detector supplied with a PC and Windows 8 or equivalent based software to control the operation of the gas chromatograph and for the acquisition and processing of data, together with a Laser printer. A library of mass spectra such as the NIST library shall also be provided to enable the identification of pesticide residues and environmental chemicals. - Column Oven - Operating temperature: ≤ 5 °C above ambient to ≥ 440 °C - Temperature set point resolution: ≤ 1 °C - Maximum temperature programming rate: >100 °C/min - Maximum run time: ≥ 900 min - Oven ramps/plateaus: 20/21,Negative ramps allowed - Autosampler: CombiPAL and Headspace sampler and holder - Injection port - Split/splitless, Multi-mode inlet with electronic pressure/flow control, PTV and others - Maximum operating temperature: at least 400°C - Pressure setting range: 0 - ≥ 6 bar - Pressure and flow settings - Pneumatics Control Unit - Electronic Pneumatics control for all gases - Inlets and Detectors should be available with Electronic Pneumatic Control on all gas flows (carrier, split vent, makeup, and combustion gases) - Pressure should be adjustable by increments of ≤ 0.02psi - Pressure/flow programming ramps: at least three - Split/splitless inlet flow sensor for control and storage of split ratio - Pneumatic Splitter: capillary flow technology for devices for effluent splitting, back-flushing and column switching - Mass selective detector - Capable of working in the MS –MS mode - Ionization mode: EI and PCI (electron impact ionization and chemical impact) - Ionization energy: Filament voltage should be user selectable over the range of 10 – 300 eV - Dual filament system for EI and single filament system for PCI - Mass range(m/z): 10 – 1,050 any subset of this range should be selectable by the user - Scan speed/rate( electronic): up to 6,250 u/sec with 0.1 amu mass resolution - Dynamic range at least 106 - Resolution (width at half height): - Selectable, 0.7 to 2.5 Daltons using default tune - Settable, 0.4 to 4.0 Daltons using custom tune - MRM speed (transitions/sec): 500 - Minimum MRM dwell: 1 Msec - Mass Filters monolithic hyperbolic gold-coated quadrupole - EI Scan Sensitivity: 12fg or less - octafluoronaphtalene(OFN) S/N ratio ≥10:1 - EI SIM Sensitivity: at least 50 fg - octafluoronaphtalene S/N ratio ≥10:1 - Mass resolution: unit mass resolution throughout the mass range - Mass axis stability: < ± 0.10 u over 24 hours(10-40oC) - Ion Source Temperature – independently heated and controlled - Quadrupole Temperature – independently heated and controlled (106-200oC) - Interface Temperature – capable of being controlled - Collision cell : Linear Hexapole - Collision cell gas: Nitrogen with Helium quenchgas for reduction of metastable helium - Collision energy : Selectable up to 60 eV - Detector: Triple-Axix HED-EM with extended-life EM and dynamically ramped-iris - Tuning: Autotune or manual tune - Total gas flow: Up to 8ml/min GCcarrier plus another 5ml/min of methane for CI operation, plus an additional 1-2 ml/min of N2 and He for the collision cell gases. - Vacuum system: Dual stage turbo-molecular pump - Automatic Liquid Sampler - Sample tray: capable of holding at least 100 sample vials, 2 ml - Injection volume range: ≤ 0.2 – ≥ 50 µl - Sampling depth should be variable - Chromatography Software - Windows compatible software for data acquisition and processing, and for control of all gas chromatograph parameters via a PC - Quantification and Calibration possibilities - Data import and export to other software - Possibility to upgrade the system for new instruments - A desk top computer suitable for running the software, equipped with Windows 8 or equivalent, high resolution LCD or LED monitor, and a laser printer (printer to be supplied with high volume cartridges sufficient for approx. 8000 to 10.000 pages). - Service/repair contract and spares for 3 years |  |  |  |
| 014 | ICP/MS - ICP/MS for analysis of elements in the full range, fully automated - Range for the mass spectrometer: 2- ≥ 260 AMU - Resolution < 0.4 AMU - Dwell time ≤ 0.1 ms - Supplied with an auto-sampler - An uninterruptible power supply (UPS) with necessary additional batteries for a minimum of 30 minutes back up. - A desk top computer suitable for running the software, equipped with Windows 8 or equivalent, high resolution LCD or LED monitor, and a laser printer (printer to be supplied with high volume cartridges sufficient for approx. 8000 to 10.000 pages). - Service/repair contract and spares for 3 years |  |  |  |
| 044 | Atomic Absorption Spectrometer with graphite furnace - Atomic absorption spectrophotometer with graphite furnace - Atomic absorption spectrophotometer with Zeeman background correction system. - Wavelength range 200 – 880 nm minimum - Computer controlled wavelength selection; automatic slit selection; - min. 4 lamp holder; automatic lamp position optimisation, - Supplied with: - Graphite Furnace - Fully computer-controlled Heated Graphite Atomizer (HGA) - Tube temperature – ambient to ≥ 2,500 °C - Supplied with a cooling system - Programmable with at least 10 independently programmable steps, parameters to include - Holding time – 0 to 99 s in steps of 1 s - Temperature ambient to ≥ 2,500 °C in steps of 10 °C - Heating rate – programmable, time dependant or maximum power - Internal gas flow - Read delay - Read time - Safety interlocks to prevent operation of the furnace: - with low gas pressure - with a broken tube - when the furnace is open - Thermal protection against overheating of the furnace - Autosampler – fully programmable and suitable for use in combination with the Graphite Furnace Atomizer - Lamps (EDL or Hollow cathode lamp) for Ag, Al, As (2 pcs), Au, Ca, Cd (2 pcs), Cr, Co, Cu (2 pcs), Fe (2 pcs), Hg, K, Mg, Mn (2 pcs), Mo, Na, Pb (2 pcs), Se, Sn, Zn (2 pcs). - Software - Windows 8 or equivalent compatible software for data acquisition and processing, and for control of all instrument parameters (including the graphite furnace and auto-sampler) - A desk top computer suitable for running the software, equipped with Windows 8 or equivalent, high resolution LCD or LED monitor, and a laser printer (printer to be supplied with high volume cartridges sufficient for approx. 8000 to 10.000 pages). |  |  |  |
| 073 | HPLC with DAD/IR detector - High Performance Liquid Chromatograph with Diode Array Detector (DAD) and Infrared Detector (IR), comprising - auto-sampler, - quaternary Gradient solvent delivery system, - DAD and IR detectors – - A desk top computer suitable for running the software - Solvent Delivery System - Quaternary pump capable of mixing up to 4 components of an eluting mixture - Flow range, capable of being set to deliver a flow rate of ≤ 0.1ml to ≥ 10ml/min with ≤ 0.1ml/min increments - Flow precision at least 0.3% RSD - Composition precision least 0.20% SD - Variable stroke volume - Maximum operating pressure at least 400 bar - Complete with appropriate on-line vacuum degasser - Autosampler-Programmable - Fitted with Rheodyne type injector with 100 µl sample loop - Injected volume – 0.1 µl to 100 µl - Precision ≤ 0.6% RSD from 5 – 100 µl, ≤ 1% from 1 – 5 µl - Sample capacity: at least 100 sample vials (2 ml) - IR Spectrometer Built-in FTIR - Wavenumber Range 4000–650 cm-1 - Data Collection Real-time, with post-run analysis available - Spectrum Type Transmittance through disc and solid-phase sample - Window Capacity 12-24 hours of chromatography - Window Temperature -100°C to +50°C - Diode Array Detector - Detector type: 1024-element diode array - Light source: Deuterium - Number of signals: 8 - Maximum sampling rate: 80 Hz (both spectra and signals) - Short-term noise < ± 3 x 10-6 - AU at 230/4 nm, slit width 4 nm, - TC 2 sec, ASTM with 10 mm Max-Light cartridge cell - Typically <± 0.6 x 10-6 - AU/cm at 230/4 nm, slit width 4 nm, - TC 2 sec, ASTM with 60 mm Max-Light cartridge cell - Drift < 0.5 x 10-3 - AU/hr at 230 nm - Linearity > 2.0 AU (5 %) at 265 nm - Typically 2.5 AU (5 %) - Wavelength range 190-640 nm - Wavelength accuracy ± 1 nm, self-calibration with deuterium lines - Slit width 4 nm (fixed) - Diode width ~ 0.5 nm - Wavelength bunching Programmable, 1 - 400 nm, in steps of 1 nm - Flow cells: - Max-Light Cartridge Cell (Standard), 10 mm, σν = 1.0 µL, 60 bar pressure maximum with RFID tags - Max-Light Cartridge Cell (High Sensitivity), 60 mm, σν = 4 µL, 60 bar pressure maximum with RFID tags - Or a detector with comparable functionalities - Chromatography Software - Windows 8 or equivalent compatible software for data acquisition and processing, and for control of all instrument parameters via a PC - Quantification and Calibration possibilities - Quality assurance system - Data import and export to other software - Possibility to upgrade the system for new instruments - A desk top computer suitable for running the software, equipped with Windows 8 or equivalent, high resolution LCD or LED monitor, and a laser printer (printer to be supplied with high volume cartridges sufficient for approx. 8000 to 10.000 pages). - Service/repair contract and spares for 3 years |  |  |  |
| 074 | HPLC with FLD and ESLD detector - High Performance Liquid Chromatograph with Fluorescence (FLD) and Evaporative Light Scattering Detector (ELSD) Detectors, comprising of: - auto-sampler, - quaternary Gradient solvent delivery system, - fluorimetric detector and a ELSD detector – - PC to control the operation of the liquid chromatograph and for the acquisition and processing of data - Solvent Delivery System - Quaternary pump capable of mixing up to 4 components of an eluting mixture - Flow range, capable of being set to deliver a flow rate of ≤ 0.1ml to ≥ 10ml/min with ≤ 0.1ml/min increments - Flow precision at least 0.3% RSD - Composition precision least 0.20% SD - Variable stroke volume - Maximum operating pressure at least 400 bar - Complete with appropriate on-line vacuum degasser - Auto-sampler-Programmable - Fitted with Rheodyne type injector with 100 µl sample loop - Injected volume – 0.1 µl to 100 µl - Precision ≤ 0.6% RSD from 5 – 100 µl, ≤ 1% from 1 – 5 µl - Sample capacity: at least 100 sample vials (2 ml) - Fluorescence detector - Excitation wavelength: from =200 nm to =700 nm and zero order, bandwidth 20nm - Emission wavelength: from =280 nm to = 900 nm and zero order, bandwidth 20nm - Sensitivity: =10 fg anthracene LOD - Standard cell volume: =10µl (standard) - Spectral acquisition: selectable for Excitation or Emission spectra - Scan speed at least 30ms per data point - Step size: selectable in the range 1 to 20 nm - Wavelength repeatability: = ±0.3nm - Wavelength accuracy: =±3nm - Evaporative Light Scattering Detector - Measurement Method Light Scattering - Light Source LED - Detection Photomultiplier Tube - Temperature Setting Range Ambient to 80°C - Nebulizer Gas Nitrogen (N2) or Air - Gas Flow Rate Max. 3.0 L/min - Gas Pressure Max. 450 kPa - Mobile Phase Flow Rate Minimum 0.2mL/min - Operating Temperature Range 5°C - 40°C - UV detector - Double beam; wavelength range 190-700nm; transmittance range 0-199.9%; concentration range -3 to 1999 ppm; - Chromatography Software - Windows 8 or equivalent compatible software for data acquisition and processing, and for control of all instrument parameters via a PC - Quantification and Calibration possibilities - Quality assurance system - Data import and export to other software - Possibility to upgrade the system for new instruments - A desk top computer suitable for running the software, equipped with Windows 8 or equivalent, high resolution LCD or LED monitor, and a laser printer (printer to be supplied with high volume cartridges sufficient for approx. 8000 to 10.000 pages). - Service/repair contract and spares for 3 years |  |  |  |
| 075 | FT-IR - For identification of organic compounds in cosmetics and drugs - Spectral range at least 7800 – 375 cm-1 - Resolution < 0.5cm-1 - RMS noise on 1 minute scan < 5.5 x10-6 Abs. - Wavelength precision at 2000 cm-1 is 0.01cm-1 or better - Spectral library for cosmetics and drugs - Auto-sampler for 20 samples or more - System should preferably be multitasking - A desk top computer suitable for running the software, equipped with Windows 8 or equivalent, high resolution LCD or LED monitor, and a laser printer (printer to be supplied with high volume cartridges sufficient for approx. 8000 to 10.000 pages). - 50 pairs of salt discs - The climate in Accra is very moist because of its position at the seaside. To protect the salt cuvettes a drying cabinet for the instrument must be provided. The instrument must be able to remain in the cabinet when being used. - Service/repair contract and spares for 3 years |  |  |  |
| 098 | Hollow cathode lamps - Hollow cathode lamps for WAGTECH S4 series AAS: - Al, As, Ca, Cd, Cr, Fe, K, one (1) each. - Cu, Hg, Pb, Sn, two (2) each. - Certified reference materials for the same metals; 5 vials per lamp. |  |  |  |
| 139 | Protein/Nitrogen analyser, Combustion method - Automated Protein/nitrogen analyser for food, feed and beverages with high capacity. - Time of analysis less than 15 minutes - Sample size > 1 g - Bases on combustion method - A desk top computer suitable for running the software, equipped with Windows 8 or equivalent, high resolution LCD or LED monitor, and a laser printer (printer to be supplied with high volume cartridges sufficient for approx. 8000 to 10.000 pages). - Service/repair contract and spares for 3 years |  |  |  |
| 184 | HPLC with UV/IR detector - High Performance Liquid Chromatograph with Ultraviolet (UV) and Infrared (IR) Detectors comprising: - auto-sampler, - quaternary Gradient solvent delivery system, - UV and IR detectors – - PC to control the operation of the liquid chromatograph and for the acquisition and processing of data - Solvent Delivery System - Quaternary pump capable of mixing up to 4 components of an eluting mixture - Flow range, capable of being set to deliver a flow rate of ≤ 0.1ml to ≥ 10ml/min with ≤ 0.1ml/min increments - Flow precision at least 0.3% RSD - Composition precision least 0.20% SD - Variable stroke volume - Maximum operating pressure at least 400 bar - Complete with appropriate on-line vacuum degasser - Autosampler-Programmable - Fitted with Rheodyne type injector with 100 µl sample loop - Injected volume – 0.1 µl to 100 µl - Precision ≤ 0.6% RSD from 5 – 100 µl, ≤ 1% from 1 – 5 µl - Sample capacity: at least 100 sample vials (2 ml) - UV / Vis Absorbance Detector - Wavelength range: ≤ 210 nm - ≥ 600 nm - Measurement range ≤ 0.01– ≥ 3.00 AUFS - Wavelength accuracy at least ±3 nm - Wavelength repeatability at least ± 0.2 nm - Automatic calibration and verification of wavelength accuracy - Complete with Standard Analytical Flow Cell - IR Spectrometer Built-in FTIR - Wavenumber Range 4000–650 cm-1 - Data Collection Real-time, with post-run analysis available - Spectrum Type Transmittance through disc and solid-phase sample - Window Capacity 12-24 hours of chromatography - Window Temperature -100°C to +50°C - Chromatography Software - Windows 8 or equivalent compatible software for data acquisition and processing, and for control of all instrument parameters via a PC - Quantification and Calibration possibilities - Quality assurance system - Data import and export to other software - Possibility to upgrade the system for new instruments - A desk top computer suitable for running the software, equipped with Windows 8 or equivalent, high resolution LCD or LED monitor, and a laser printer (printer to be supplied with high volume cartridges sufficient for approx. 8000 to 10.000 pages). - Service/repair contract and spares for 3 years |  |  |  |

Requested training:

| Spec | Equipment | Training |
| --- | --- | --- |
| 009 | HPLC MS/MS | 5 days for max. 4 persons |
| 010 | Gas chromatograph – triple quadrupole mass spectrometer | 4 days for max. 4 persons for each of the concerned laboratories |
| 014 | ICP/MS | 3 days for max. 4 persons |
| 044 | Atomic Absorption Spectrometer with graphite furnace | 4 days for max. 4 persons |
| 073 | HPLC with DAD/IR detector | 3 days for max. 4 persons |
| 074 | HPLC with FLD and ESLD detector | 3 days for max. 4 persons for each of the concerned laboratories |
| 075 | FT-IR | 1 day for max. 4 persons |
| 184 | HPLC with UV/IR detector | 3 days for max. 4 persons |

General Remarks:

- If calibration is required, the calibration certificate must be issued by a calibration laboratory accredited by an accreditation body, which is a signatory of the Mutual Recognition Arrangement of the International Laboratory Accreditation Cooperation (ILAC-MRA).

- All electrical equipment should be suitable for operating on 220-240V / 50-60 Hz (unless otherwise specified), and provided with a BS-1363 plug, or with a suitable adapter.

- Software that is required for the realisation of functionalities of equipment must be the latest version, and should be compatible with a Windows 8 or equivalent operating system. Software updates should be provided for free during at least for a period of 36 months.

- For computers/system controllers and related supplementary equipment (e.g. flat screen, printer, UPS) that are considered to be an integrated part of the equipment configuration and that are not further specified, the contractor will provide systems of good quality and reputable brands, based on latest technology, and latest software for operating systems (Windows 8, or equivalent), for office applications (MS Office, or equivalent), and anti-virus protection at least for a period of 36 months.