# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

Contract title: Supply of laboratory and metrology equipment for testing and calibration laboratories in Ghana

# Publication reference: EuropeAid/133510/D/SUP/GH

Column 1-2 should be completed by the Contracting Authority

Column 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

LOT 3: MASS MEASUREMENT

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 001 | Balance, Analytical, 4 place, ≥ 200g - Capacity ≥ 200 g - Readability ≤ 0,1 mg - Repeatability ≤ ± 0,1 mg - Linearity ≤ ± 0,2 mg - Response time ≤ 3s - Pan diameter > 80 mm - Capable of being tared across the whole weighing range - LED Display - Calibration certificate from an accredited calibration laboratory. |  |  |  |
| 004 | Balance, Top Pan – 3 place, 400g - Capacity ≥ 400 g - Single Range - Readability 0,001 g - Repeatability ≤ ± 0,001 g - Linearity ≤ ± 0,002 g - Response time ≤ 2 s - Pan diameter 110 - 160 mm - Capable of being tared across the whole weighing range - LED Display - Calibration certificate from an accredited calibration laboratory. |  |  |  |
| 005 | Weights – for monitoring balances - Stainless steel weights manufactured according to the OIML recommendations - Range 1g to 100 g - Individuals weights – minimum requirement - 1 x 1g, 1 x 5g, 1 x 10g, 1 x 20g, 1 x 50g and 1 x 100g - without an accredited calibration certificate - in a purpose made storage box - with suitable forceps |  |  |  |
| 038 | Balance, Top pan, 600g, 3 place - Capacity ≥ 600 g - Single Range - Readability 0,001 g - Repeatability ≤ ± 0,001 g - Linearity ≤ ± 0,002 g - Response time ≤ 2 s - Pan diameter 110 - 160 mm - Capable of being tared across the whole weighing range - LED Display - Calibration certificate from an accredited calibration laboratory. |  |  |  |
| 042 | Moisture analyser, laboratory - Laboratory moisture analyser for food and feed products - Rapid analysis, based on "drying on the scale" principle - Max sample size >50g - Repeatability < 0.3% at 2g |  |  |  |
| 045 | Thermogravimetry - Temperature data - Small furnace (SF) - Temperature range >… 900 ºC - Temperature accuracy 1) ± 0.5 K or better - Temperature precision 1) ± 0.15 K - Furnace temperature resolution 0.02 K - Heating time 5 min (RT … 900 ºC) - Cooling time 20 min (900 … 100 ºC) - Heating rate 2) 250 K/min - Sample volume ≤100 μl - A desk top computer suitable for running the software, equipped with Windows 8 or equivalent, high resolution LCD or LED monitor, and a laser printer (printer to be supplied with high volume cartridges sufficient for approx. 8000 to 10.000 pages). - Service/repair contract and spares for 3 years |  |  |  |
| 054 | Balance, Analytical, 4 place, ≥ 300g - Capacity ≥ 300 g - Readability ≤ 0,1 mg - Repeatability ≤ ± 0,1 mg - Linearity ≤ ± 0,2 mg - Response time ≤ 3s - Pan diameter > 80 mm - Capable of being tared across the whole weighing range - LED Display - Calibration certificate from an accredited calibration laboratory. |  |  |  |
| 055 | Balance, Top pan, 1600 g - Top Pan Balance - Capacity: > 1 600 g - Readability: 0.01 g - Tare facility - Digital Display - Stainless Steel Weighing pan: rectangular - Calibration certificate from an accredited calibration laboratory. |  |  |  |
| 056 | Sliding triple beam balance, 2000g - Capacity > 2000g - Readability ≤1g - Calibration certificate from an accredited calibration laboratory. |  |  |  |
| 088 | Balance, Top Pan – 2 place, 400g - Capacity ≥ 400 g - Single Range - Readability 0,01 g - Repeatability ≤ ± 0,01 g - Linearity ≤ ± 0002 g - Response time ≤ 2 s - Pan diameter 110 - 160 mm - Capable of being tared across the whole weighing range - LED Display - Calibration certificate from an accredited calibration laboratory. |  |  |  |
| 134 | Balance, Top Pan – 1 place, 4000g - Top Pan Balance - Capacity: > 4 000 g - Readability: 0.1 g - Tare facility - Digital Display - Stainless steel weighing pan: rectangular, approx dimensions 200 mm x 200mm. - Calibration certificate from an accredited calibration laboratory. |  |  |  |
| 141 | Balance, Top Pan – 1 place, 3000g - Top Pan Balance - Capacity: > 3 000 g - Readability: 0.1 g - Tare facility - Digital Display - Stainless steel weighing pan: rectangular, approx dimensions 200 mm x 200mm - Calibration certificate from an accredited calibration laboratory |  |  |  |
| 200 | Set of Weights, OIML Class F1, 1g to 1 kg - Range: 1g; 2 g; 2g; 5g; 10g; 20g; 20g; 50g; 100g; 200g; 200g; 500g; 1000g. - Without an adjustment chamber. - With a knob. - With a wooden box for storage. - Calibration certificate from an accredited calibration laboratory |  |  |  |
| 344 | Set of Weights OIML class F1, 20 kg (Quantity: 50) and 1000 kg (Quantity: 2) - Features: - Rectangular block weights with fixed handle/lifting ring - Material: 20 kg rectangular block weight: High Grade Stainless Steel, 1000 kg: Stainless steel or other metal of equivalent hardness (outer material) - Specifications: - Class F1 as per OIML R-111-1 - Average density: 7390 – 8730 kgm-3 - Including: - Calibration certificate from an accredited calibration laboratory for all weights - User manual in English |  |  |  |
| 345 | 2000 kg Mass comparator F2 (500 kg to 2000 kg) - Features: - Manual mass comparator - Suitable for comparison of class F2 weights from 500 kg to 2000 kg - Continuous electrical weighing range - Manual/crane loading of the weights - Automated weighing process and data storage - Filters can be optimally adapted to the particular ambient conditions - Separate electronic unit to reduce the effect of heat generated during operation - Built-in application programs: For mass determination (ABA and ABBA cycles), density determination, time controlled functions, differential weighing, statistics, air buoyancy correction, etc. - Control/Data port to PC - Specifications: - Maximum capacity: 2100 kg - Application range: 0-2100 kg - Readability: 1 g - Repeatability, s: 8 µg - Repeatability (typical), s: 5 µg - Electronic taring range: 2100 kg - Linearity: 30 g - Off-centre loading error: 0.7 g / mm - Stabilization time: 30 s - Including: - Data cable and mass calibration software - System controller PC - Calibration certificate from an accredited laboratory - On-site installation and testing - Complete technical documentation in English including operating, service and programming manuals - Service/repair contract and spares for 3 years |  |  |  |
| 372 | Weights set 1 g to 20 kg OIML class F1 in box, with accessories - Features: - Material: High grade stainless steel - Values: 1-2-2-5 sequence - Specifications: - class F1 in accordance with OIML R-111-1 - Including: - Strong wooden storage box - Tweezers, forceps and gloves for handling weights - Cleaning/maintenance accessories - Calibration certificate from an accredited calibration laboratory for all weights - User manual in English |  |  |  |
| 373 | Digital scale, capacity 60 kg - Features: - Portable - Bench or platform type - Operation on rechargeable battery, at least 48 hours on one charge - Specifications: - Class II in accordance with OIML R-76-1 - Dual range: Max = 30 kg, d = 1 g, e = 5 g and Max = 60 kg, d = 1 g, e = 10 g - Tare range 30 kg - Including: - Spare battery - Charger/adapter - System controller PC (laptop) - Calibration certificate from an accredited calibration laboratory User manual in English - Extended warranty plus fast delivery of spare parts |  |  |  |
| 375 | Test weights 1000 kg for weighbridges, class M2 - Features: - 40 test weights for weighbridge verification - Construction: cylindrical with handle for movement by rolling - Material: Cast iron and steel assembly, suitably painted, in accordance with OIML R-111-1, average density > 3000 kg/m3 - Specifications: - Class M2 in accordance with OIML R-111-1 - Tolerance in mass: ±160 g - Including: - Calibration certificate from an accredited calibration laboratory for all weights - Any relevant accessories |  |  |  |
| 381 | Digital Crane scale, 2000 kg, accuracy 0.1% - Features: - Robust, weather proof measuring unit be attached between the crane hook and the load - Large digital display - Wireless communication to handheld remote readout unit (included) or laptop PC. Range: (open-air, line-of-sight) 60 m - Data storage and analysis functions - Specifications: - Resolution: 2 kg and 0.5 kg selectable - Safe overload capacity: 200 % of nominal load - Uncertainty: 0.1% of nominal capacity - Repeatability: 0.1% of capacity - Display update rate: 2 times per second - Battery life: 320 hours typical use with two C-cell alkaline batteries, 40 hours typical with use of wireless communication to remote readout unit - Operating temperature: 0 °C to 60 °C - Including: - Remote readout unit with communication port to PC for transfer of measurement data; OR alternatively: communication hardware and any required software drivers to laptop PC for data transfer directly to PC (laptop PC included fully installed with relevant software) - Data analysis software to generate measurement reports - Calibration certificate from an accredited calibration laboratory - User manual in English - Extended warranty plus fast delivery of spare parts |  |  |  |
| 393 | Set of Weights OIML Class F1, 1 mg to 10 kg - Features: - Weights sequence: 1-2-2-5 - Material: High grade stainless steel (1 g and higher) - Specifications: - Class F1 as per OIML R-111-1 - Density: 7390 – 8730 kgm-3 - Including: - Calibration certificate from an accredited calibration laboratory for all weights - Tweezers, forceps and gloves for handling weights - Robust wooden storage box and maintenance accessories - User manual in English |  |  |  |
| 394 | Single weight F1, 2 g - Features: - Material: High grade stainless steel - Specifications: - Class F1 as per OIML R-111-1 - Average density: 7390 – 8730 kgm-3 - Including: - Calibration certificate from an accredited calibration laboratory |  |  |  |
| 395 | 100 g Manual Mass Comparator F1 (1 mg to 100 g) - Features: - Suitable for comparison of class F1 weights from 1 mg to 100 g - Continuous electrical weighing range - Manual loading of the weights - Automated weighing process and data storage - Filters can be optimally adapted to the particular ambient conditions - Separate electronic unit to reduce the effect of heat generated during operation - Built-in application programs: For mass determination (ABA and ABBA cycles), density determination, time controlled functions, differential weighing, statistics, air buoyancy correction, etc. - Motorized draft shield with conductive coating to prevent the effects of static electricity - Control/Data port to PC - Specifications: - Maximum capacity: 111 g - Application range: 0-111 g - Readability: 1 µg - Repeatability, at 1/10 load: 1 µg - Repeatability, s: 2 µg - Repeatability (typical), s: 1 µg - Electronic taring range: 61 g - Linearity: 10 µg - Off-centre loading error: 2 µg / mm - Range sensitivity: 2 µg / 10 mg - Stabilization time: 15 s - Including: - External calibration weight, 50 g E2 - Data cable and mass calibration software - Calibration certificate from an accredited laboratory - On-site installation and testing - Complete technical documentation in English including operating, service and programming manuals - Service/repair contract and spares for 3 years |  |  |  |
| 396 | 2 kg Manual Mass Comparator F1 (200 g – 2 kg) - Features: - Suitable for comparison of class F1 weights from 200 g to 2 kg - Continuous electrical weighing range - Manual loading of the weights - Automated weighing process and data storage - Filters can be optimally adapted to the particular ambient conditions - Separate electronic unit to reduce the effect of heat generated during operation - Built-in application programs: For mass determination (ABA and ABBA cycles), density determination, time controlled functions, differential weighing, statistics, air buoyancy correction, etc. - Motorized draft shield with conductive coating to prevent the effects of static electricity - Control/Data port to PC - Specifications: - Maximum capacity: 2500 g - Application range: 0-2500 g - Readability: 100 µg - Repeatability, at 1/10 load: 100 µg - Repeatability, s: 200 µg - Repeatability (typical), s: 100 µg - Electronic taring range: 2500 g - Linearity: 1 mg - Off-centre loading error: 30 µg / mm - Range sensitivity: 0.1 mg / 2 g - Stabilization time: 10 s - Including: - External calibration weight, 2 kg E2 - Data cable and mass calibration software - Calibration certificate from an accredited calibration laboratory - On-site installation and testing - Complete technical documentation in English including operating, service and programming manuals - Service/repair contract and spares for 3 years |  |  |  |
| 397 | 60 kg Manual Mass Comparator F1 (5 kg – 50 kg) - Features: - Suitable for comparison of class F1 weights from 5 kg – 50 kg - Continuous electrical weighing range - Manual loading of the weights - Automated weighing process and data storage - Filters can be optimally adapted to the particular ambient conditions - Separate electronic unit to reduce the effect of heat generated during operation - Built-in application programs: For mass determination (ABA and ABBA cycles), density determination, time controlled functions, differential weighing, statistics, air buoyancy correction, etc. - Control/Data port to PC - Specifications: - Maximum capacity: 64 kg - Application range: 0-2500 g - Readability: 2 mg - Repeatability, at 1/10 load: 4 mg - Repeatability, s: 6 mg - Repeatability (typical), s: 4 mg - Electronic taring range: 2500 g - Linearity: 70 mg - Off-centre loading error: 3.5 mg / mm - Range sensitivity: 20 mg / 100 g - Stabilization time: 10 s - Including: - External calibration weight, 50 kg E2 - Data cable and mass calibration software - Calibration certificate from an accredited calibration laboratory - On-site installation and testing - Complete technical documentation in English including operating, service and programming manuals - Service/repair contract and spares for 3 years |  |  |  |
| 417 | System controller PC, Lot 3 - Desk top computer - Processor: Intel-i7, or equivalent, or better - Internal memory: at least 8 Gb - Hard drive: at least 500 Gb, HHD and/or SSD - Keyboard and wireless mouse - Colour monitor at least 20" flat widescreen with zero dead pixels guarantee (ISO 13406-2 Class II for 3 years) - Communication: WIFI, BT, at least 3 USB slots, and HDMI. Webcam, speakers, and microphone integrated - Windows 8 or equivalent - Software for word-processing, spreadsheets, e-mail, presentation, and internet browser, compatible with the latest Microsoft Office programmes (2013 version) - Virus protection and internet security software with subscription for 3 years - Supplied with a UPS for 15 minutes |  |  |  |

Requested training:

| Spec | Equipment | Training |
| --- | --- | --- |
| 045 | Thermogravimetry | 2 days for max. 4 persons for each of the concerned laboratories |
| 344 345 | Set of weights F1 2000 kg mass comparator F2 | 1 day for 3 persons |
| 381 | Digital crane scale (to be combined with Lot 8 / spec 380) | 1 day for max. 4 persons |
| 395 396 397 | Mass comparator 100g Mass comparator 2 kg Mass comparator 60 kg | 1 day (for all specs combined) for max. 4 persons |

General Remarks:

- If calibration is required, the calibration certificate must be issued by a calibration laboratory accredited by an accreditation body, which is a signatory of the Mutual Recognition Arrangement of the International Laboratory Accreditation Cooperation (ILAC-MRA).

- All electrical equipment should be suitable for operating on 220-240V / 50-60 Hz (unless otherwise specified), and provided with a BS-1363 plug, or with a suitable adapter.

- Software that is required for the realisation of functionalities of equipment must be the latest version, and should be compatible with a Windows 8 or equivalent operating system. Software updates should be provided for free during at least for a period of 36 months.

- For computers/system controllers and related supplementary equipment (e.g. flat screen, printer, UPS) that are considered to be an integrated part of the equipment configuration and that are not further specified, the contractor will provide systems of good quality and reputable brands, based on latest technology, and latest software for operating systems (Windows 8, or equivalent), for office applications (MS Office, or equivalent), and anti-virus protection at least for a period of 36 months.