# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

Contract title: Supply of laboratory and metrology equipment for testing and calibration laboratories in Ghana

# Publication reference: EuropeAid/133510/D/SUP/GH

Column 1-2 should be completed by the Contracting Authority

Column 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

LOT 4: SIMPLE ANALYSERS

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 017 | pH Meter, laboratory - pH Range: ≤ 0.00 to ≥ 14.00 - Resolution: pH – 0.01 - Accuracy: ≤ ± 0.02 pH - Temperature Compensation: Automatic or manual - Automatic buffer recognition - Calibration: Automatic / manual override - 3 point calibration with possibility of using 1 or 2 points - Supplied with: - Electrode stand - 2 Glass electrodes incorporating a pH sensing membrane and internal Ag/AgCl reference electrode - Temperature compensation probe. - Reference solutions, traceable to international recognised references: - pH 4: 5 x 1 l - pH 7: 5 x 1 l - pH 10: 5 x 1 l |  |  |  |
| 025 | Polarimeter, Semiautomatic - Semi-automatic Polarimeter complete with replacement tubes (100 and 200m) - Accuracy: Angle of rotation: ≤ ±0.1° - Range: Rotation: ~ +180° to ~ - 180° - Resolution: Angle: ≤ 0.05° - Tube length: up to 200 mm - Supplied with a quartz reference plate |  |  |  |
| 027 | Butyrometer - Set of butyrometers for: - Milk: 0 – 8 % - Milk: 0 – 10 % - Cream: 0 – 70 % - Cheese: 0 – 40 % - Skimmed milk: 0.5 % - 10 pieces of each - Each with a calibration certificate from an accredited calibration laboratory |  |  |  |
| 029 | Combustion calorimeter - Automated combustion calorimeter for determination of energy content in food, drinks, and feed - Liquid and solid samples - Automatic oxygen filling, decomposition, and sample ignition - If cooling is needed a closed cooling system must be supplied - Service/repair contract and spares for 3 years |  |  |  |
| 030 | Fibre analysis system - Fibre analyse system for determination of crude fibres and dietary fibres in food and feed products - Capacity: more than 25 samples per 8 hours - Automated system |  |  |  |
| 036 | Spectrophotometer - UV/Vis - Range < 200 to > 900 nm - Double beam - Wavelength accuracy: ≤ 1 nm - Bandwidth ≤ 4 nm - Mode options of absorbance, transmittance or concentration - Photometric range ≤ -0.3 to ≥ 2.5 A, 0 to ≥ 130 % - T, 0 – 9 999 C - Deuterium lamp, Tungsten Halogen lamp - Automatic cell changer with 4 or more cells - To be supplied with: - spare lamps, one each - Quartz cuvettes 10mm, 10 pcs. - Glass cuvettes 10mm, 10 pcs. - Teflon covers, 20 pcs. - A printer – internal or external - Calibration certificate from an accredited calibration laboratory. |  |  |  |
| 040 | Protein/Nitrogen analyser, Kjeldahl principle - Automated Protein/nitrogen analyser based on the Kjeldahl principle for food, feed and beverages with high capacity. - With an auto-sampler of 20 or more spaces - Time of analysis less than 15 minutes - Sample size > 1 g - A desk top computer suitable for running the software, equipped with Windows 8 or equivalent, high resolution LCD or LED monitor - Service/repair contract and spares for 3 years |  |  |  |
| 043 | Refractometer, high precision - For refractive index measurements on a wide variety of materials including sugars, confectionery, pharmaceuticals, fats and oils - Enclosed glass scale calibrated to 0.001 R.I. with estimation to 0.0001. - Hinged prism box with water jacket and built-in temperature probe - Measuring temperature range: 10°C below ambient to > 99°C - Temperature accuracy < ± 0.05 °C - Built-in digital temperature readout with LCD - LED scale illuminator - Low voltage external power supply - Flow cell accessory - To be connected to the water tap - Refractive index < 1.30 to > 1.74 - Range 0.0 - 95.0% sugar - Calibration certificate from an accredited calibration laboratory. |  |  |  |
| 061 | Flame photometer - Flame Photometer for Na, K, Ca, Ba, and Li - With filters or grids for the above elements - Uncertainty < 0.5 % - Any gas will be supplied by the beneficiary |  |  |  |
| 072 | Total fat analysing system, fully automated - Fully automated fat analysing system based on the Soxhlet Extraction System for analysing fat in food and feed - The system must be able to operate overnight analysis - The instrument must provide a means for the automatic, accelerated, Soxhlet extraction process. - Multiple extraction programs should be stored and available for different sample types and solvents. Up to six extractions shall be performed simultaneously. - Multiple heating programs to describe variable heating levels and time settings. - Microprocessor controlled temperature monitoring and control with excess temperature protection. Continuous monitoring of cooling water with automatic shut-off. - Temperature range: up to 280 °C or more - Acid vapours generated during the digestion must be removed safely without any danger to the operator. - The equipment should include the extraction unit with control unit, 100 extraction thimbles, 6 holders for extraction thimbles, insert rack, 12 extraction beakers, boiling chips, compressor, O-rings and all connections. - Service/repair contract and spares for 3 years |  |  |  |
| 091 | Sulphur analyser for fuels - Analyser for sulphur in petroleum products according to ASTM D4294 - Automated system for at least 10 samples - Range 0.003 - 6 wt%, repeatability 10 ppm |  |  |  |
| 092 | Melting point apparatus - Automated system for determining melting point of chemical substances. - Temperature range: Ambient to >360°C - Temperature resolution: 0.1°C - Temperature accuracy: 0.5°C at 20°C and 1.5°C at 360°C - Digital display |  |  |  |
| 096 | Auto-sampler for WAGTECH AAS - Auto-sampler for WAGTECH S series atomic absorption spectrometer for the flame mode - Has to be used with WAGTECH S4 AAS Serial number GE712063 |  |  |  |
| 099 | Dairy product analyser - Automated system for fast analysis of dairy products - The following parameters must be included: fat, protein, lactose, total solids - The system must be supplied with an auto sampler with at least 50 positions. - Service/repair contract and spares for 3 years |  |  |  |
| 100 | pH Meter, hand held - For field work - pH Range: ≤ 0.00 to ≥ 14.00 - Resolution: < 0.02 pH - Accuracy: ≤ ± 0.05 pH - Reference solutions, traceable to international recognised references: - pH 4: 5 x 1 l - pH 7: 5 x 1 l - pH 10: 5 x 1 l |  |  |  |
| 133 | Auto-dilutor for WAGTECH AAS - Auto-dilutor for WAGTECH S4 series atomic absorption spectrometer for the flame mode - Has to be used with WAGTECH S4 AAS Serial number GE712063 |  |  |  |
| 151 | Soil Texture Apparatus - Soil texture apparatus according to the USDA Method - Sieve Shaker- 01 pc capable of - Memory for 10 test sieve parameter sets - Timer function: 0 - 99 min or continuous operation - Intermittent or continuous shaking operation - Digital display of all functions, Timer, Speed, - Variable speed control 0-1000 rpm - Maximum loading: 3 kg per sieve - A set of standard sieves 20 cm diameter opening 2, 1, 0.5, 0.25 and 0.125mm, lid and bottom pan, 2 sets - Sample divider: mechanical as per USDA, 1 pc - Soil Stirrer Disperser, 1 pc - volumes from 0.25 - 30 l (H2O) - speed range from 500 – 10,000 rpm - Digital speed display - Electronic speed control - Electronic overload protection - Stainless steel dispersing elements suitable for soil dispersing - Bottle top Dispenser 50ml - Replaceable dispensing cartridge and valves - Easy cleaning and maintenance - Easy to operate - Refills automatically - Volume range from 0.2 to 50 ml - ASTM 152H Hydrometers- 10pcs |  |  |  |
| 152 | Guelph Permeameter - Guelph Permeameter(- ASTM D5126) consisting of - permeameter - field tripod - well auger - well preparation and cleanup tools - collapsible water container - vacuum test hand pump |  |  |  |
| 153 | Tensiometer - For measurement of soil moisture potential - Heavy walled (minimum 6.4mm) durable butyrate tubing 7/8 in. (2.22 cm) OD with screw cap to allow for easy replacement of the ceramic cup, o-ring sealed dial gauge, and use of addition of extension tubes. - Readout gauges hermetically sealed in a stainless steel case with vent access to assure precision readings. - tensiometer lengths, (45 cm) Insertion Tools can be used for coring a hole in the soil to accept these units. - Accessories include - Service Kit - tensiometer filling fixture compatible with the tensiometer - tensiometer insertion tools minimum 300 mm, 600 mm,900 mm, and 1200 mm - Recalibrator vacuum dial |  |  |  |
| 154 | Soil Salinity meter, Portable - Portable Soil Salinity meter. The meter measures the soil conductivity in EC (mS/cm) as well as salt activity (g/L). To measure directly the soil salinity range 0-100 ohms. Uses common battery or mains - Features - Probes for liquid and direct soil measurement - Automatic temperature compensation - BEPS (Battery Error Prevention System) alerts the user in the event that low battery power could adversely affect readings - Includes probes, battery, instructions, and rugged carrying case |  |  |  |
| 155 | Soil moisture meter, Portable - Soil moisture content measurement from 0 to > 50% - Min/Max records minimum and maximum moisture readings - Data hold to freeze reading on display - Water resistant housing - Replaceable Heavy Duty Moisture Probe - Complete with 4 x AAA batteries and sensor cap - Calibration certificate from an accredited calibration laboratory |  |  |  |
| 156 | Penetrometer, Field - Features include: - Large, easy-to-read display - Maximum pressure 100KN/m2 - Multiple display units: lbs; psi; psf; tsf, or kg per cm2 - Dual-rod design, which eliminates soil friction |  |  |  |
| 157 | Penetrometor, Pocket - Indicates consistency, shear strength, and approximate unconfined shear strength. - Direct-reading scale—in tons/sq ft, or kg/sq cm—corresponds to equivalent unconfined compressive strength. - Range: 0 to 4.5 tons. - High quality construction. |  |  |  |
| 165 | DNA based Pathogen detection kits - DNA based pathogen detection kits including - Equipment for sample preparation - PCR Cycler for microbiological identification - Complete detection system with electrophoresis system for microbiological use - Gel imaging system with high resolution camera including software and computer - A desk top computer suitable for running the software, equipped with Windows 8 or equivalent, high resolution LCD or LED monitor - Software must be compatible with Windows 8 or equivalent - Tools, Pipettes, Disposables to run the system - Reagents and protocols for - Campylobacter, - yeast & mould, - Vibrio, - S. aureus, - Salmonella assays, - Listeria assays, - Enterobacter sakazakii, - E. Coli - Service/repair contract and spares for 3 years |  |  |  |
| 174 | Auto-analyser for food samples - Auto-analyser for food and environment samples - Discrete analyser - Capacity: more than 50 samples per hour - Reading positions: minimum 50 measuring (cuvette) positions, - Heated positions, 37°C: Minimum 50, - Kinetic and end point measurements, - Software complying with - An easy-to-use user interface, - Ability to run batch-serially or random access. - Extensive QC protocols are available, which includes password protection, data integrity, data safety and audit trails. - A desk top computer suitable for running the software, equipped with Windows 8 or equivalent, high resolution LCD or LED monitor, and a laser printer (printer to be supplied with high volume cartridges sufficient for approx. 8000 to 10.000 pages). - Expandable Discrete Chemistry Analyser configured to analyse the following parameters in food and environment samples: - Acetaldehydes, - total acidity, - amino acids, - ascorbic acid - benzoic acid, - caffeine, - carbohydrates, - citric acids, - Chlorides, - Chlorine, - Cholesterol, - Fructose, - Galactose, - Glycerol, - Hydroxymethylfurfural, - Iodine, - Reducing sugars, - Starch, - Sucrose, - Sulphur dioxide, - Methanol, - Colour |  |  |  |
| 194 | Seed counter - Automatic seed counter for seeds size 0.3 - 15 mm. - Counts more than 1.000.000 seeds. - Complete with a desk top computer suitable for running the software, equipped with Windows 8 or equivalent, high resolution LCD or LED monitor, and a laser printer (printer to be supplied with high volume cartridges sufficient for approx. 8000 to 10.000 pages). |  |  |  |
| 195 | Seed counter, vacuum - Pneumatic sucking, 50 – 100 seeds, push system to release seeds on petri dish, - minimum 1,200 watts, - brass/acrylic heads 50 – 100 holes |  |  |  |

Requested training:

| Spec | Equipment | Training |
| --- | --- | --- |
| 029 | Combustion calorimeter | 1 day for max. 4 persons |
| 030 | Fibre analysis system | 2 days for max. 4 persons |
| 040 | Protein/Nitrogen analyser, Kjeldahl principle | 2 days for max. 4 persons |
| 072 | Total fat analysing system, fully automated | 1 day for max. 4 persons for each of the concerned laboratories |
| 099 | Dairy product analyser | 2 days for max. 4 persons |
| 139 | Protein/Nitrogen analyser, Combustion method | 2 days for max. 4 persons |
| 165 | DNA based pathogen detection kits | 3 days for max. 4 persons |
| 174 | Auto-analyser for food samples | 4 days for max. 4 persons |

General Remarks:

- If calibration is required, the calibration certificate must be issued by a calibration laboratory accredited by an accreditation body, which is a signatory of the Mutual Recognition Arrangement of the International Laboratory Accreditation Cooperation (ILAC-MRA).

- All electrical equipment should be suitable for operating on 220-240V / 50-60 Hz (unless otherwise specified), and provided with a BS-1363 plug, or with a suitable adapter.

- Software that is required for the realisation of functionalities of equipment must be the latest version, and should be compatible with a Windows 8 or equivalent operating system. Software updates should be provided for free during at least for a period of 36 months.

- For computers/system controllers and related supplementary equipment (e.g. flat screen, printer, UPS) that are considered to be an integrated part of the equipment configuration and that are not further specified, the contractor will provide systems of good quality and reputable brands, based on latest technology, and latest software for operating systems (Windows 8, or equivalent), for office applications (MS Office, or equivalent), and anti-virus protection at least for a period of 36 months.