# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

Contract title: Supply of laboratory and metrology equipment for testing and calibration laboratories in Ghana

# Publication reference: EuropeAid/133510/D/SUP/GH

Column 1-2 should be completed by the Contracting Authority

Column 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

LOT 5: PHYSICAL TESTING EQUIPMENT

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 124 | Plastic pipe testing machine - Suitable for testing all kinds of liquid transportation with Plastic Pipes and Pipe Static Hydraulic Blasting Trail (including U-series PVC Pipes, PE Pipes, PE Gas Pipes, Pressure vessels etc.) - Simultaneously on multiple tube specimens under different pressure test, should not interfere with each other. - Automatic alarm and automatically shut off when sample raptures. - Pressure Range: 0.6 MPa – 10MPa (any setting) - Pressure Resolution: ±0.01MPa - Pressure Control Precision: +2%, -1% - Timing Show Range: 0 – 999h - Timing Precision: 1% - Timing range: 1 min – 1000h - Timing Accuracy: ±0.1% - Number of sample:5 - Dimension of the Test Piece: Suitable for diameter Ø14 – Ø800 mm - Testing tank: - Inside Dimension: Approx 1500 x 1000 x 1300 mm - Temperature control range: ambient – 95°C - Constant Temperature Accuracy: ±1°C - Temperature Gradient: ±1°C - Jigs (end caps) in pairs - Ø32, Ø40, Ø50, Ø63, Ø75, Ø90, Ø110, Ø160, Ø225, Ø315, & Ø400 mm |  |  |  |
| 125 | Tensile strength testing machine - Test/Jog Speeds: 0.001 – 500 mm per minute at force up to 25 KN and to 250 mm per minute up to the full 50KN capacity of the load frame. - Return speeds: 0.001 – 500mm per minute. - Test Clearances: - Vertical clearance to 1100 mm with any tooling and a horizontal clearance of 405 mm between the columns, approximately. |  |  |  |
| 126 | Pensky-Martens Closed Cup Flash Point Apparatus - An automated flash point tester capable in fulfilling the test method ASTM-D93. - This must come with a sample changer for at least 8(eight) test cups. - Automatically controlled ignition, stirring speed and flash detection. - Working temperature range: ambient to 400°C. - The cooling /cool down time between tests should not exceed 10 (ten) minutes - There should be an in-built safety measure for fire and over-heating. |  |  |  |
| 127 | Cleveland Open Cup Flash Point Apparatus - An automated flash point tester capable of fulfilling the test method ASTM-D92. - Automatic controlled stirrer, ignition and flash detector - Operating temperature range from ambient to 400°C. - The cooling /cool down time between tests should not exceed 10 (ten) minutes - There should be an in-built safety measure for fire and over-heating. |  |  |  |
| 128 | Distillation unit for testing fuels - Stand-alone unit, fully automated in operation and complying with ASTMD86 and ASTMD1160. - Independent temperature control of distillation flask, condenser and receiver. - It should not use refrigerants / coolants but with a Peltier cell cooling. |  |  |  |
| 129 | Reference materials for flash point testers - Certified reference material for flash point testers with certificates - Flash points: ranges: - 45 - 55°C, Accuracy: ±2°C - 95 - 105°C, Accuracy: ±2°C - 140 - 160°C, Accuracy: ±5°C - 180 - 220°C, Accuracy: ±5°C |  |  |  |

Requested training:

| Spec | Equipment | Training |
| --- | --- | --- |
| 124 | Plastic pipe testing machine | 2 days for max. 4 persons |

General Remarks:

- If calibration is required, the calibration certificate must be issued by a calibration laboratory accredited by an accreditation body, which is a signatory of the Mutual Recognition Arrangement of the International Laboratory Accreditation Cooperation (ILAC-MRA).

- All electrical equipment should be suitable for operating on 220-240V / 50-60 Hz (unless otherwise specified), and provided with a BS-1363 plug, or with a suitable adapter.

- Software that is required for the realisation of functionalities of equipment must be the latest version, and should be compatible with a Windows 8 or equivalent operating system. Software updates should be provided for free during at least for a period of 36 months.

- For computers/system controllers and related supplementary equipment (e.g. flat screen, printer, UPS) that are considered to be an integrated part of the equipment configuration and that are not further specified, the contractor will provide systems of good quality and reputable brands, based on latest technology, and latest software for operating systems (Windows 8, or equivalent), for office applications (MS Office, or equivalent), and anti-virus protection at least for a period of 36 months.