# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

Contract title: Supply of laboratory and metrology equipment for testing and calibration laboratories in Ghana

# Publication reference: EuropeAid/133510/D/SUP/GH

Column 1-2 should be completed by the Contracting Authority

Column 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

LOT 6: DIMENSIONAL MEASUREMENT

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 301 | Laser length standard with gauge block interferometer - Features: - For calibration of K and 00 grades gauge blocks - Two frequency stabilized laser light sources - Measures gauge length, flatness and length variation over the measuring faces of the gauge block - Automatic measurement and compensation for ambient conditions - Automatic calculation and print-out of measurement results, selectable inch or metric units - Specifications: - Maximum length of gauge block: 300 mm - Measurement uncertainty: (0.05 + 0.3 x L) µm (L in meters) - Lowest uncertainty attainable: 0.02 µm - Including: - Software for measurement and automatic reporting of rectangular, square gauge blocks and round length bars, with uncertainty analysis - Installation and extensive testing - Complete with accessories, line stabiliser / UPS, system controller PC and software - Complete documentation including installation test report, operation and service manuals in English - Service contract and spares for 3 years |  |  |  |
| 302 | Grade K stainless steel gauge blocks - Features: - 122 pieces: 1.005 mm to 100 mm - Specifications: - Length tolerances: - 0.5-10 mm: 0.20 µm - 10-25mm: 0.30 µm - 25-50mm: 0.40 µm - 50-75mm: 0.50 µm - 75-100mm: 0.60 µm - Including: - Calibration certificate from an accredited calibration laboratory - Strong wooden storage box - Maintenance kit |  |  |  |
| 303 | Grade 00 stainless steel gauge blocks - Features: - 122 pieces: 1.005 mm to 100 mm - Specifications: - Length tolerances: - 0.5-10 mm: 0.05 µm - 10-25mm: 0.05 µm - 25-50mm: 0.10 µm - 50-75mm: 0.13 µm - 75-100mm: 0.15 µm - Including: - Calibration certificate from an accredited calibration laboratory - Strong wooden storage box - Maintenance kit |  |  |  |
| 304 | Gauge block comparator for length up to 250 mm - Features: - Dual gauge heads - Semi automatic with temperature sensors - Specifications: - Resolution: 0.01 µm, uncertainty: (0.03 + 0.3 x L) µm (L in meters), or better - Including: - Gauge block set and holder for probe calibration - software, for determination of central length, minimum length, maximum length and parallelism - Complete with full set of accessories - Complete operation and service manuals in English - Extended warranty plus fast delivery of spare parts |  |  |  |
| 305 | Manual height Gauge block comparator for length 250 mm up to 750 mm - Features: - For rectangular, circular, square gauge blocks and length bars - Dual gauge heads - Specifications - Resolution: 0.01 µm - Including: - Gauge block set and holder for probe calibration - Complete with full set of accessories - Complete operation and service manuals in English - Extended warranty plus fast delivery of spare parts |  |  |  |
| 306 | Precision Vertical Height gauge, 0 to 1000 mm - Features: - Carbide-tipped scribes - Display: Zero setting, Pre-setting, Probe tip diameter compensation, data hold, data output in mm or inch - SPC data output - Specifications: - Resolution: 0.001 mm and 0.0001” selectable - Repeatability: 0.001 mm or better - Including: - System controller PC and software - Complete with full set of accessories - Complete operation and service manuals in English - Extended warranty plus fast delivery of spare parts |  |  |  |
| 307 | Hardware and software upgrade for existing universal length machine model TRIMOS LABCONCEPT 1000 - Supply of: - Adapters: Clamping attachment for ring gauges up to 60 mm (ref no. 14.2) - Measuring device for ring gauges 1 to 100 mm (ref no. LABC 70) - Set of ruby ball inserts (ref no. TEL 76) - Set of T-shaped ruby ball inserts (ref no. TEL 75) - Dial indicator holding device (ref no, TULM 13.2) - Measuring arm for interchangeable probes (ref no. TELMA 16) - Floating measuring table (ref no. TELS 6) - Thread wire sets 0.25 to 5.5 mm pitch - Dial indicator to fit on the dial indicator holding device - Software upgrade (ref no. QM MANAG 32) and module for measuring device (LABC 70) - Double USB interface for connection of the measurement units to the controller - Work included: - Installation and testing of the upgraded assembly - Include: - Calibration certificates from an accredited calibration laboratory for any relevant upgrade parts delivered - Operation manuals in English for adapters and software |  |  |  |
| 308 | Calibration bench for line measures up to 2 m and tape measures up to 10 m - Features: - Horizontal bench with tape loading facility and cross-hair microscope on travelling carriage - Measuring length: 2 m - Digital display readout and data output to controller - Technology: linear encoder or laser interferometer - Temperature sensors in the bench in at least 5 positions - Specifications: - Resolution: 1 µm - Displacement uncertainty: 5 µm or better over 2 m length - Temperature uncertainty: 0.2 °C - Including: - System controller PC and software - Complete with full set of accessories - Calibration certificates from an accredited calibration laboratory - Complete operation and service manuals in English - Service contract and spares for 3 years |  |  |  |
| 309 | Angle standards - Features: - Sine bar with accessories - Made of hardened steel - With detachable end stop - For use with gauge blocks (not included) - Set of 12 angle standards: 15’, 30’, 1°, 2°, 3°, 4°, 5°, 10°, 15°, 20°, 25° and 30° - Made of hardened steel - Precision Square - Made of hardened steel - Specifications: - Sine bar - Angle range: 0 to 60° - Circularity error of sine bar cylinders: 0.2 μm - Centre distance: 127 mm ± 0.004 mm - Angle standards: - Angle uncertainty: 0,001° (4 arc seconds) - Width: 6,35 mm, Length: 76,2 mm - Precision Square (flat) - Beam lengths/Accuracy at full length: 150 mm x 100 mm/ 0.008 mm and 250 mm x 165 mm/0.010 mm - Including: - Carrying cases, maintenance kits and any other necessary accessories - Calibration certificates from an accredited calibration laboratory - Operation manuals in English |  |  |  |
| 310 | Optical flats/parallel sets - Features: - Ranges: 0-25 mm, 25-50mm, 0-1” and 1-2” - 4 thicknesses in each set - Specifications: - Diameter: 30 mm - Flatness: 0.1 µm and 0.000004” respectively - Including: - Carrying case, maintenance kit and any other necessary accessories - Calibration certificates from an accredited calibration laboratory - Operation manuals in English |  |  |  |
| 311 | Precision stainless steel measuring tapes 10 m, 50 m and 100 m - Features: - Tape thickness: 0.25 mm maximum - With holder and winding mechanism - Specifications: - Uncertainty: 0.05 mm per meter length or better - Including: - Calibration certificates from an accredited calibration laboratory |  |  |  |
| 312 | Granite surface plates - Features: - Sizes: 1.5 x 1.2 x 0.3 m and 0.4 x 0.4 x 0.05 m - Specifications: - Flatness/surface grade: 0 - Including: - Stand for the large surface plate (with integrated cupboard/drawers) - Calibration certificates from an accredited calibration laboratory |  |  |  |

Requested training:

| Spec | Equipment | Training |
| --- | --- | --- |
| 301 | Laser length standard with gauge block interferometer | 3 days for max. 2 persons |
| 302 303 304 305 306 310 | Gauge block grade K Gauge blocks grade 00 Gauge block comparator Manual height gauge block comparator Precision vertical height gauge Optical flats | 1 day (combined for all specs) for max. 2 persons |
| 307 | Update universal length machine | 1 day for max. 2 persons |
| 308 311 309 | Tape and line measure calibration bench Precision measuring tapes Angle standards | 1 day (combined for all specs) for max. 2 persons |

General Remarks:

- If calibration is required, the calibration certificate must be issued by a calibration laboratory accredited by an accreditation body, which is a signatory of the Mutual Recognition Arrangement of the International Laboratory Accreditation Cooperation (ILAC-MRA).

- All electrical equipment should be suitable for operating on 220-240V / 50-60 Hz (unless otherwise specified), and provided with a BS-1363 plug, or with a suitable adapter.

- Software that is required for the realisation of functionalities of equipment must be the latest version, and should be compatible with a Windows 8 or equivalent operating system. Software updates should be provided for free during at least for a period of 36 months.

- For computers/system controllers and related supplementary equipment (e.g. flat screen, printer, UPS) that are considered to be an integrated part of the equipment configuration and that are not further specified, the contractor will provide systems of good quality and reputable brands, based on latest technology, and latest software for operating systems (Windows 8, or equivalent), for office applications (MS Office, or equivalent), and anti-virus protection at least for a period of 36 months.