# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

Contract title: Supply of laboratory and metrology equipment for testing and calibration laboratories in Ghana

# Publication reference: EuropeAid/133510/D/SUP/GH

Column 1-2 should be completed by the Contracting Authority

Column 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

LOT 8: FORCE, TORQUE, AND PRESSURE

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 341 | Electronic Force Transducers 10 kN, 50 kN, 200 kN, 1000 kN and 5000 kN with dual input amplifier / indicator - Features: - For calibrating force machines and force transducers by comparison - 6-wire transducer connections, sensor type: DMS-full bridge - Control/Data interface to PC - Specifications: - Transducers: - Uncertainty 0.05 % of range (typical) - Dual input amplifier/indicator: - Accuracy class: 25 ppm - Nonlinearity: ±5 ppm - Reproducibility: ±5 ppm - Temperature drift: Zero point: < ±1 ppm/K, sensitivity: < ±2,5 ppm/K - Drift: < ±5 ppm (after calibration) - Long term stability: < ±25 ppm/a - Bridge resistance: 150…5000 Ω at 5V, 300 … 5000 Ω at 10V, 1000…5000 Ω at 20V - Input signal Range: ±40 mV - Maximum input voltage: 10 V - Excitation voltage: 5, 10 or 20 VDC - Integration time electronics: 0.1 to 5000 ms - Integration time software: programmable moving average filter - Maximum sampling rate: 50 readings per second - Maximum resolution: ±200 000 - Noise (3σ-value) at 500 ms: ±10 ppm - Including: - All cables and connectors as required, including data cables to controller PC - Any relevant accessories - Control and data analysis software for generating reports - Calibration certificates from an accredited calibration laboratory for each component - User manual in English - Service/repair contract and spares for 3 years |  |  |  |
| 342 | Universal force machine up to 6000 kN - Features: - For generating compression force for calibrating force transducers by comparison - Specifications: - Stability: at least sufficient for use with items 341 at specified accuracy - Including: - Data cables to controller - Software drivers for control of the machine - Calibration certificates from an accredited calibration laboratory for each component - User manual in English - Service/repair contract and spares for 3 years |  |  |  |
| 343 | Torque transducers (standards) ranges 50 Nm, 200 Nm, 1000 Nm and 5000 Nm (class 0.05%) with dual input amplifier/indicator, accessories, controller hardware and software - Features: - For calibrating torque machines and torque transducers by comparison - Specifications: - Typical uncertainty: 0.05% of range - Including: - All cables and connectors as required, including data cables to controller - Data analysis software for generating reports - Calibration certificates from an accredited calibration laboratory for each component - User manual in English - Service/repair contract and spares for 3 years |  |  |  |
| 414 | Torque generator for calibration of torque transducers up to 5000 Nm - Features: - For generating torque for calibrating torque transducers by comparison - Specifications: - Stability: at least sufficient for use with items 343 at specified accuracy - Including: - Data cables to controller - Software drivers for control of the machine - Calibration certificates from an accredited calibration laboratory for each component - User manual in English - Service/repair contract and spares for 3 years |  |  |  |
| 346 | High pressure primary oil balance 10 to 5000 bar - Features: - Measuring Range 10 – 5000 bar gauge - More than one piston cylinder assembly with interchangeable cylinders - Material of construction: Whole tungsten carbide for Piston assemblies - Trim mass set 10mg - Individual mass value ≤ 5kg - Free deformation Piston cylinder Mounting system - Built-in power rotation of piston - Built-in spirit level and levelling feet - Built-in 4 wires platinum resistance thermometer - Built-in pressure generation and control - Active electronics in separate control unit - Specifications: - Uncertainty of masses ≤ 10 ppm, calibrated to a gravity acceleration (g) of 9.78109 ± 1.10-7 ms-2 - Area Stability: ≤ 1 ppm/year - Resolution: ≤ 1x10-6 FS - Typical uncertainty (k =2): 50 ppm of reading - Including: - Pressure Gauge 0 to 5 bar - All connectors, adaptors and tubing needed for connecting various equipment under test - Data port and connection cable to controller - System controller PC - Data analysis software for generating calibration certificates - Calibration certificates from an accredited calibration laboratory - User manual in English - Service/repair contract and spares for 3 years |  |  |  |
| 347 | Deadweight tester 8000 bar - Features: - Complete system including weights set - Specifications: - Uncertainty: 0.05%, calibrated to a gravity acceleration (g) of 9.78109 ± 1.10-7 ms-2 - Including: - All connectors, adaptors and tubing needed for connecting various equipment under test - Calibration certificate from an accredited calibration laboratory - Data interface and cabling to controller - Software for control and measurement data analysis - User manual in English - Service/repair contract and spares for 3 years |  |  |  |
| 348 | Absolute pressure balance from 0 to 2 bar (abs), incl. vacuum pump - Features: - Range 0 – 2 bar Absolute - Technology based on a Piston-Cylinder and a Force Cell - Internal vacuum chamber as reference chamber - Built-in Environmental Monitoring Module to do all the compensation from environmental influences - Connection capability to an automatic pressure controller - The Pressure Balance and Controller (System) has to have the capability to be controlled from a PC - Unit to have self calibrating facility - Specifications: - Resolution ≤ 1x10-6 FS - Typical uncertainty (k =2) U= 0.8 Pa +1x10-5 P, calibrated to a gravity acceleration (g) of 9.78109 ± 1x10-7 ms-2 - Including: - Vacuum pump 5 m3/h - All connectors, adaptors and tubing needed for connecting various equipment under test - Data port and connection cable to controller - System controller PC - Control and data analysis software for generating calibration certificates - Calibration certificate from an accredited calibration laboratory - User manual in English - Service/repair contract and spares for 3 years |  |  |  |
| 349 | Pneumatic pressure controller -1 to +5 bar - Features: - Range -1 bar to +5 bar - Specifications: - Accuracy ≤ ± 0.01 % FS - Including: - All connectors, adaptors and tubing needed for connecting various equipment under test - Data port and connection cable to controller - Data analysis and control software - Calibration certificate from an accredited calibration laboratory - User manual in English - Service/repair contract and spares for 3 years |  |  |  |
| 350 | Primary Pure gas Pressure Balance 0.2 to 10 bar - Features: - Operation on pure gas pneumatic pressure up to 50 bar input pressure - External pressure control through a precision control block - One piston-cylinder mounting system - Pressure range 0.2 to 10 bar - Material of construction whole tungsten carbide with titanium head - Re-entrant mounting system - Built-in power rotation of piston - Built-in spirit level and levelling feet - Built-in platinum resistance thermometer - Built-in inductive piston position detectors - Built-in pressure generation and control - No active electronics in piston assembly - Must be easily upgradeable to higher pressure range - Specifications: - Area stability: ≤ 1 ppm/year - Resolution: ≤ 0.2 Pa - Typical uncertainty (k =2): 15 ppm of reading, calibrated to a gravity acceleration (g) of 9.78109 ± 1.10-7 ms-2 - Including: - Manual pressure control block up to 200 bar input pressure - Pressure gauge 0 to 10 bar - All connectors, adaptors and tubing needed for connecting various equipment under test - Data port and connection cable to controller - Data analysis software for generating calibration certificates - Calibration certificate from an accredited calibration laboratory - User manual in English - Service/repair contract and spares for 3 years |  |  |  |
| 351 | Pressure Balance for Differential Pressure and Primary Gas Pressure Balance 5 to 800 bar (abs), with booster pump - Features: - Differential Pressure balance with two piston-cylinders mounting systems - Static Gauge pressure 5 to 800 bar - Differential pressure 0 to 195 bar, with P + DP < 800 bar - Material of construction: whole tungsten carbide for piston assemblies - Operating Principle: piston-cylinder assembly must be of the liquid lubricated type - Mounting system: re-entrant - Built-in power rotation of piston - Built-in spirit level and levelling feet - Built-in 4 wire platinum resistance thermometer - Built-in inductive piston position detectors - Built-in pressure generation and control - No active electronics - Must be easily upgradeable to higher pressure range - Specifications: - Uncertainty of masses: ≤ 10 ppm, calibrated to a gravity acceleration (g) of 9.78109 ± 1.10-7 ms-2 - Trim mass set 10 mg - Individual mass value: ≤ 2 kg - Area Stability: ≤ 1 ppm/year - Resolution: ≤ 10 Pa - Typical uncertainty (k =2): 20 ppm of reading - Including: - Gas pressure intensifier up to 1000 bar - All connectors, adaptors and tubing needed for connecting various equipment under test - Data port and connection cable to controller - Data analysis software for generating calibration certificates - Calibration certificate from an accredited calibration laboratory - User manual in English - Service/repair contract and spares for 3 years |  |  |  |
| 352 | Oil-gas separator repair/upgrade - Description: - Replacement of existing tubing and connectors by steel parts (supply of parts and materials included). Full refurbishment works of the complete system. Testing up to maximum working pressure. - Oil/gas separator is available and in working condition. - Specifications: - System must be suitable for working pressure up to 400 bar - Including: - Final test report |  |  |  |
| 353 | Working standard 0 to 2000 bar (hydraulic) - Features: - Manual oil pressure generator with deadweight pressure balance - Weights made of non-magnetic, austenitic, stainless steel - Specifications: - Uncertainty: 0.02 % of reading (from 10 % to 100 % of range) , calibrated to a gravity acceleration (g) of 9.78109 ± 1.10-7 ms-2 - Including: - All connectors, adaptors and tubing needed for connecting various equipment under test - Data analysis software for generating calibration certificates - Calibration certificate from an accredited calibration laboratory - User manual in English - Extended warranty plus fast delivery of spare parts |  |  |  |
| 354 | Pressure controller with digital reference gauge and small gas bottle, 200 bar - Features: - Complete portable system for on-site calibrations: Gas cylinder up to 200 bar, manual pressure controller 1 to 200 bar, digital pressure indicator - Specifications: - Uncertainty: 0.1 % of full scale - Temperature coefficient (for gauge pressure): < 0.02 % of reading per °C (from 0 °C to 50 °C) - Including: - Small size gas bottle, about 5 L or smaller for 200 bar compressed air, with valve and tubing to pressure controller - All connectors, adaptors and tubing needed for connecting various equipment under test - Data port and connection cable to controller - Data analysis software for generating calibration certificates - Calibration certificate from an accredited calibration laboratory - User manual in English - Extended warranty plus fast delivery of spare parts |  |  |  |
| 355 | Digital reference gauge and small gas bottle, 200 bar - Description: - The same specification as item 354, but without the pressure controller |  |  |  |
| 356 | Pressure chamber for barometer calibration - Features: - Pressure chamber with pressure controller - Inner dimensions: about 60 x 40 x 40 cm or more - Computer control of pressure - Built-in reference digital pressure indicator - Pressure range: 0.7 to 1.3 bar - Specifications: - Pressure stability: 0.1 mbar - Control/gauge accuracy: 0.1 mbar - Including: - Data port and connection cable to controller - Control and data analysis software for generating calibration certificates - Calibration certificate from an accredited calibration laboratory - User manual in English - Extended warranty plus fast delivery of spare parts |  |  |  |
| 357 | Digital Pressure Indicator with manual pump for field work up to 700 bar (water) - Features: - Complete portable system for on-site calibrations - Manual pump up to 700 bar (using distilled water) - Digital pressure indicator as pressure reference - Specifications: - Uncertainty (gauge pressure): 0.025 % of full scale - Including: - Adapters for various quick connect and basic thread fittings - Data port and connection cable to PC - Data analysis software for generating calibration certificates - Calibration certificate from an accredited calibration laboratory - User manual in English - Extended warranty plus fast delivery of spare parts |  |  |  |
| 377 | Tire pressure gauge calibrator - Features: - Digital pressure calibrator with integrated pump, up to 16 bar or more, operated on rechargeable battery - Fitted with connection and adapters for various types of tire pressure gauges - Indicates gauge pressure in bar - Specifications: - Pressure range: 0 to 16 bar or more - Resolution: 0.01 bar - Maximum deviation, including calibration uncertainty: 0.02 bar in range from 0 to 4 bar and 0.16 bar in range from 4 to 10 bar and over the temperature range from 15 to 40 °C - Battery life: at least 48 hours operation on one charge - Including: - Carrying case - Any relevant accessories - Charger/adapter - Calibration certificate from an accredited calibration laboratory - User manual in English - Extended warranty plus fast delivery of spare parts |  |  |  |
| 380 | Electronic Force Transducers 10 kN, 50 kN, 200 kN, 1000 kN and 5000 kN with Amplifier/indicator - Features: - For calibrating force machines and force transducers by comparison - 6-wire transducer connections, sensor type: DMS-full bridge - Control/Data interface to PC - Specifications: - Transducers: - Uncertainty 0.1 % of range - Dual input amplifier/indicator: - Accuracy class: 100 ppm of range - Bridge resistance: 150…5000 Ω at 5V, 300…5000 Ω at 10V, 1000…5000 Ω at 20V - Input signal Range: ±40 mV - Maximum input voltage: 10 V - Excitation voltage: 5, 10 or 20 VDC - Integration time electronics: 0.1 to 5000 ms - Integration time software: programmable moving average filter - Maximum sampling rate: 50 readings per second - Including: - All cables and connectors as required, including data cables to controller PC - Any relevant accessories - System controller PC (laptop) - Control and data analysis software for generating reports - Calibration certificate from an accredited calibration laboratory for each component - User manual in English - Extended warranty plus fast delivery of spare parts |  |  |  |
| 398 | Dead weight tester (pneumatic) 0 to 200 bar - Features: - Manual pump and external supply from gas bottle - Precision piston gauge with weights, calibrated to a gravity acceleration (g) of 9.78109 ± 1.10-7 ms-2 - Specifications: - Uncertainty: 0.02 % of value or better - Including: - Tubing, adaptors and connectors for various types of equipment under test - Robust storage case for weights - Calibration certificate from an accredited calibration laboratory for piston gauge and weights - User manuals in English - Extended warranty and fast delivery of spare parts |  |  |  |
| 399 | Dead weight tester (hydraulic) 0 to 700 bar - Features: - Oil pressure bench top system - Pressure generation: internal manual pump - Precision piston gauge with weights, calibrated to a gravity acceleration (g) of 9.78109 ± 1.10-7 ms-2 - Specifications: - Uncertainty: 0.02 % of value or better - Including: - Tubing, adaptors and connectors for various types of equipment under test - Robust storage case for weights - Calibration certificate from an accredited calibration laboratory for piston gauge and weights - User manuals in English - Extended warranty and fast delivery of spare parts |  |  |  |
| 400 | Portable pressure calibrators pneumatic 0 to 200 bar - Features: - Complete portable calibration system - Pressure generation: Hand pump for lower range to 40 bar and input port for external supply from gas bottle for range up to 200 bar - Needle valve for accurate pressure control - Venting valve - System must be capable of maintaining a stable pressure - Test and reference ports - Operating temperature (ambient): 0 to +50°C - Digital reference pressure indicator to 200 bar - Operating time on batteries: 30 hours or more - Over pressure protection valve - Sensor input for comparison of unit under test with reference - Data storage: minimum 300 measurements - Data port to PC - Specifications: - Uncertainty: reference pressure indicator: 0.1 % of reading - Over pressure range: 220 bar - Including: - All tubing, connectors and adaptors required for various types of pressure gauges under test - Small gas bottle 5 L or less for 200 bar air - Data cable to PC and analysis software for generating calibration reports - Calibration certificate from an accredited calibration laboratory for reference pressure indicator - User manuals in English - Extended warranty and fast delivery of spare parts |  |  |  |
| 401 | Portable pressure calibrators hydraulic, 0 to 700 bar - Features: - Complete portable calibration system - Pressure generation: Hydraulic hand pump - Needle valve for accurate pressure control - Test and reference ports - Digital reference pressure indicator 0 to 700 bar - Over pressure protection valve - Pressure range: 0 to 700 Bar - Pressure media: Oil - Operating time on batteries: Approximately 35 hours - Maximum weight: 9 kg - Operating temperature (ambient): 0 to +50°C - Process temperature: -20 to +80°C - Sensor input for comparison of unit under test with reference - Data storage: 30 sensors, 10 points each - Functions: Programmable Input: mA/mV/V/Switch Test, Range, Conversion, Engineering Units, Filter, Min/Max/Tare, History Log, Internal/External Sensor, Battery Indicator, Leak Test - Data interface to PC - Specifications: - Accuracy ± 0,03 % FS. (10-40 °C), temperature effects: 0,001 % of reading per °C - mA input 0-30 mA (± 0,03 % FS), V input 0-15 VDC (± 0,05 % FS), at ambient temperature 25°C, temperature effects: ± 0,008 %/°C - Pressure overload: FS x 1.5 - Including: - All tubing, connectors and adaptors required for various types of pressure gauges under test - Data cable to PC and analysis software for generating calibration reports - Calibration certificate from an accredited calibration laboratory for reference pressure indicator - User manuals in English - Extended warranty and fast delivery of spare parts |  |  |  |
| 410 | System controller PC, Lot 8 - Desk top computer - Processor: Intel-i7, or equivalent, or better - Internal memory: at least 8 Gb - Hard drive: at least 500 Gb, HHD and/or SSD - Keyboard and wireless mouse - Colour monitor at least 20" flat widescreen with zero dead pixels guarantee (ISO 13406-2 Class II for 3 years) - Communication: WIFI, BT, at least 3 USB slots, and HDMI. Webcam, speakers, and microphone integrated - Windows 8 or equivalent - Software for word-processing, spreadsheets, e-mail, presentation, and internet browser, compatible with the latest Microsoft Office programmes (2013 version) - Virus protection and internet security software with subscription for 3 years - Supplied with a UPS for 15 minutes |  |  |  |
| 415 | Set of IEEE-488 (GPIB) cables - Shielded with stackable male/female GPIB connectors at both ends - Qty and length: 1 x 2 m, 2 x 1 m and 3 x 0.5 m |  |  |  |
| 416 | IEEE-488 to USB adapter - For connecting instruments with IEEE-488 (GPIB) bus to a computer (PC or Laptop) - Including hardware drivers for Windows operating systems and com drivers for MS Excel and other applications |  |  |  |
| 425 | Precision barometric pressure monitor - Features: - Precision barometer (absolute pressure meter) with - Large digital display - Selectable units including mmHg, bar and mbar/hectopascal - Inbuilt data logger function of at least 50 000 measurement values at selectable intervals from 10 seconds to 1 hour - Data transfer to computer - Operation on rechargeable battery and line power, alternatively non-rechargeable battery operation with battery life of 4 years or more - Specifications: - Pressure range: at least including 750-1150 mbar absolute - Accuracy: 0.4 mbar over temperature range of at least 15 °C to 45 °C - Including: - Battery charger/adaptor in case of rechargeable battery - Spare rechargeable battery - Data cable/interface to PC - Data analysis/reporting software for PC/Windows - Calibration certificate from an accredited calibration laboratory - Extended warranty plus fast delivery of spare parts |  |  |  |

Requested training:

| Spec | Equipment | Training |
| --- | --- | --- |
| 346 347 348 349 350 351 352 353 354 355 356 357 | High pressure primary oil balance Dead weight tester Absolute pressure balance Pneumatic pressure controller Primary pure gas pressure balance Pressure balance for differential pressure Oil gas separator Working standard Pressure controller Digital reference gauge Pressure chamber for barometer calibration Digital pressure indicator | 4 days (all specs combined) for 2 persons |
| 377 | Tire pressure gauge calibrator | ½ day for max. 4 persons |
| 380 | Electronic force transducers (to be combined with Lot 3 / spec 381) | 1 day for max. 4 persons |
| 398 399 400 401 082 415 416 410 | Dead weight tester Dead weight tester Portable pressure calibrators Portable pressure calibrators Environmental monitor T/RH IEEE-488 cable set IEEE-488 to USB adapter System controller | 1 day (all specs combined) for max. 4 persons |

General Remarks:

- If calibration is required, the calibration certificate must be issued by a calibration laboratory accredited by an accreditation body, which is a signatory of the Mutual Recognition Arrangement of the International Laboratory Accreditation Cooperation (ILAC-MRA).

- All electrical equipment should be suitable for operating on 220-240V / 50-60 Hz (unless otherwise specified), and provided with a BS-1363 plug, or with a suitable adapter.

- Software that is required for the realisation of functionalities of equipment must be the latest version, and should be compatible with a Windows 8 or equivalent operating system. Software updates should be provided for free during at least for a period of 36 months.

- For computers/system controllers and related supplementary equipment (e.g. flat screen, printer, UPS) that are considered to be an integrated part of the equipment configuration and that are not further specified, the contractor will provide systems of good quality and reputable brands, based on latest technology, and latest software for operating systems (Windows 8, or equivalent), for office applications (MS Office, or equivalent), and anti-virus protection at least for a period of 36 months.

Additional information on selected items for better understanding the nature of the repairs / refurbishment required

| Spec | Equipment | Description |
| --- | --- | --- |
| 352 | Oil-gas separator repair/upgrade | - Existing equipment: - The Gas/Oil Separator is of Brand/Type: Hensel / HE MTG 400.01 - The existing system consists of a standard gas cylinder outside of the laboratory and tubing valves and pressure gauges, essentially as shown in the figure below. - Condition of existing equipment: - The existing tubing is unsuitable for the pressures used, but was used because of unavailability of the correct type. - Works and supplies included in this item: - The works under this tender include the replacement of this tubing, connectors and valves. - The existing Oil/Gas separator should be maintained but pressure gauges, valves and reduction valves should be replaced with new ones. - Checking of the system for leaks and testing of proper operation are part of the included work and a test report should be made, before the commissioning of the system. - The length of tubing from the outdoor gas cylinder to the pressure reduction valve inside the laboratory is approximately 100 metres maximum |

Schematic drawing of the installation: